## Remenant
XYZ MAPPIN | TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Mon, Feb 09, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
0AMPS ON ALL COILS
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
Step | Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt | Time
0 | 0 | 0 | 0 | 0.57 | -1.446 | -21.16 | 21.217 | 15:24
1 | 1 | 0 | 0 | 0.72 | -1.191 | 6.34 | 6.491 | 15:24
2 | 10 | 0 | 0 | 0.89 | -1.164 | 13.35 | 13.4302 | 15:24
3 | 20 | 0 | 0 | 0.94 | -1.067 | 22.38 | 22.4251 | 15:25
4 | 30 | 0 | 0 | 0.98 | -0.955 | 31.35 | 31.3798 | 15:25
5 | 40 | 0 | 0 | 0.89 | -0.854 | 38.62 | 38.6397 | 15:25
6 | 50 | 0 | 0 | 0.7 | -0.792 | 42.3 | 42.3132 | 15:26
7 | 60 | 0 | 0 | 0.42 | -0.797 | 41.64 | 41.6497 | 15:26
8 | 70 | 0 | 0 | 0.14 | -0.859 | 36.37 | 36.3804 | 15:26
9 | 80 | 0 | 0 | -0.07 | -0.984 | 27.41 | 27.4277 | 15:26
10 | 100 | 0 | 0 | -0.25 | -1.28 | 4.61 | 4.7909 | 15:27
11 | 110 | 0 | 0 | -0.17 | -1.407 | -5.49 | 5.67 | 15:27
12 | 120 | 0 | 0 | -0.02 | -1.465 | -13.12 | 13.2016 | 15:27
13 | 130 | 0 | 0 | 0.19 | -1.484 | -17.99 | 18.0521 | 15:28
14 | 140 | 0 | 0 | 0.35 | -1.485 | -20.37 | 20.4271 | 15:28
15 | 150 | 0 | 0 | 0.46 | -1.467 | -21.25 | 21.3055 | 15:28
16 | 160 | 0 | 0 | 0.53 | -1.441 | -21.34 | 21.3952 | 15:28
17 | 170 | 0 | 0 | 0.57 | -1.423 | -21.08 | 21.1357 | 15:29
18 | 180 | 0 | 0 | 0.58 | -1.433 | -20.92 | 20.977 | 15:29
19 | 190 | 0 | 0 | 0.56 | -1.438 | -20.99 | 21.0467 | 15:29
20 | 200 | 0 | 0 | 0.55 | -1.449 | -21.26 | 21.3164 | 15:30
21 | 210 | 0 | 0 | 0.53 | -1.441 | -21.51 | 21.5647 | 15:30
22 | 220 | 0 | 0 | 0.54 | -1.43 | -21.22 | 21.275 | 15:30
23 | 230 | 0 | 0 | 0.57 | -1.375 | -19.69 | 19.7462 | 15:31
24 | 240 | 0 | 0 | 0.66 | -1.268 | -16.01 | 16.0737 | 15:31
25 | 250 | 0 | 0 | 0.75 | -1.088 | -9.82 | 9.9085 | 15:31
26 | 260 | 0 | 0 | 0.86 | -0.856 | -0.72 | 1.4109 | 15:32
27 | 270 | 0 | 0 | 0.96 | -0.627 | 10.61 | 10.6718 | 15:32
28 | 280 | 0 | 0 | 1 | -0.437 | 22.43 | 22.4565 | 15:32
29 | 290 | 0 | 0 | 0.94 | -0.336 | 32.84 | 32.8552 | 15:32
30 | 300 | 0 | 0 | 0.79 | -0.346 | 40.19 | 40.1993 | 15:33
31 | 310 | 0 | 0 | 0.61 | -0.497 | 43.25 | 43.2572 | 15:33
32 | 320 | 0 | 0 | 0.4 | -0.747 | 41.69 | 41.6986 | 15:33
33 | 330 | 0 | 0 | 0.17 | -1.041 | 35.91 | 35.9255 | 15:34
34 | 340 | 0 | 0 | 0.01 | -1.315 | 27.41 | 27.4415 | 15:34
35 | 350 | 0 | 0 | -0.03 | -1.507 | 18.03 | 18.0929 | 15:34
36 | 360 | 0 | 0 | 0.03 | -1.608 | 9.46 | 9.5957 | 15:34
37 | 370 | 0 | 0 | 0.12 | -1.616 | 2.52 | 2.996 | 15:35
38 | 380 | 0 | 0 | 0.24 | -1.579 | -2.59 | 3.0429 | 15:35
39 | 390 | 0 | 0 | 0.36 | -1.509 | -6.05 | 6.2457 | 15:35
40 | 400 | 0 | 0 | 0.44 | -1.442 | -8.05 | 8.19 | 15:36
41 | 0 | 0 | 0 | 0.64 | -0.995 | -19.79 | 19.8253 | 15:36
41 | 0 | 0 | 0 | 0.44 | -1.054 | 13.28 | 13.329 | 15:36

## Main bore alignment II
XYZ MAPPIN | TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Mon, Feb 09, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
CUT 10 POINT "WAFERS" AT BORE, 1/3 DOWN Z AXIS, CENTER, AND 2/3 DOWN Z AXIS THEN OTHER BORE.
THIS TEST DONE WITH 3 AXIS HALL, TO ENSURE MA | NETIC FIELD IS ALI | NED WITH BORE
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
Step | Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt | Time
1 | 20 | -20 | 0 | -1722.5 | -14.55 | -11029.5 | 11163.2024 | 14:36
2 | 20 | -10 | 0 | -941.7 | -19.09 | -11112.6 | 11152.4456 | 14:37
3 | 20 | 0 | 0 | -173.5 | -23.44 | -11145.7 | 11147.075 | 14:37
4 | 20 | 10 | 0 | 589.9 | -27.29 | -11137.2 | 11152.845 | 14:37
5 | 20 | 20 | 0 | 1363.9 | -30.19 | -11084.7 | 11168.3351 | 14:38
6 | 20 | 0 | 20 | -153 | -333.28 | -10943.7 | 10949.8427 | 14:38
7 | 20 | 0 | 10 | -159.7 | -176.97 | -11065.2 | 11067.7673 | 14:38
8 | 20 | 0 | 0 | -171.7 | -24.03 | -11144.5 | 11145.8485 | 14:39
9 | 20 | 0 | -10 | -179.6 | 126.73 | -11176.3 | 11178.4614 | 14:39
10 | 20 | 0 | -20 | -189.6 | 281.4 | -11166.2 | 11171.3543 | 14:39
11 | 67.5 | -20 | 0 | -1624.9 | -37.68 | -18236.5 | 18308.7862 | 14:39
12 | 67.5 | -10 | 0 | -920 | -40.72 | -17999 | 18022.5431 | 14:40
13 | 67.5 | 0 | 0 | -240.4 | -44.31 | -17915.5 | 17917.1676 | 14:40
14 | 67.5 | 10 | 0 | 430 | -47.92 | -17974.6 | 17979.8065 | 14:40
15 | 67.5 | 20 | 0 | 1113.6 | -51.47 | -18180.5 | 18214.6462 | 14:41
16 | 67.5 | 0 | 20 | -218.8 | -281.22 | -18151.5 | 18154.9968 | 14:41
17 | 67.5 | 0 | 10 | -226.5 | -161.73 | -17956.1 | 17958.2568 | 14:41
18 | 67.5 | 0 | 0 | -237.9 | -45.29 | -17914 | 17915.6368 | 14:41
19 | 67.5 | 0 | -10 | -249.7 | 68.68 | -18014.4 | 18016.2614 | 14:42
20 | 67.5 | 0 | -20 | -264.4 | 183 | -18266.1 | 18268.9301 | 14:42
21 | 199.5 | -20 | 0 | -186.7 | -50 | -19964.5 | 19965.4356 | 14:42
22 | 199.5 | -10 | 0 | -188.2 | -49.42 | -19963.6 | 19964.5482 | 14:43
23 | 199.5 | 0 | 0 | -186.8 | -49.3 | -19963 | 19963.9348 | 14:43
24 | 199.5 | 10 | 0 | -188 | -49.1 | -19963 | 19963.9456 | 14:43
25 | 199.5 | 20 | 0 | -189.2 | -48.83 | -19963.2 | 19964.1563 | 14:43
26 | 199.5 | 0 | 20 | -186.7 | -49.17 | -19965.2 | 19966.1335 | 14:44
27 | 199.5 | 0 | 10 | -187.4 | -49.11 | -19964.1 | 19965.0399 | 14:44
28 | 199.5 | 0 | 0 | -186.7 | -49.35 | -19963.2 | 19964.134 | 14:44
29 | 199.5 | 0 | -10 | -186.6 | -49.51 | -19963.1 | 19964.0335 | 14:45
30 | 199.5 | 0 | -20 | -184.7 | -49.68 | -19963.8 | 19964.7162 | 14:45
31 | 331.5 | -20 | 0 | 1642.3 | -8.17 | -13367.3 | 13467.8107 | 14:45
32 | 331.5 | -10 | 0 | 770.4 | -5.12 | -13328.5 | 13350.7473 | 14:46
33 | 331.5 | 0 | 0 | -55.3 | -0.52 | -13309.9 | 13310.0149 | 14:46
34 | 331.5 | 10 | 0 | -875.3 | 4.55 | -13311 | 13339.7486 | 14:46
35 | 331.5 | 20 | 0 | -1727.2 | 9.98 | -13330 | 13441.4367 | 14:46
36 | 331.5 | 0 | 20 | -84.4 | 330.79 | -13477.7 | 13482.0229 | 14:47
37 | 331.5 | 0 | 10 | -68.7 | 161.34 | -13383.5 | 13384.6488 | 14:47
38 | 331.5 | 0 | 0 | -57.9 | 0.87 | -13311.3 | 13311.426 | 14:47
39 | 331.5 | 0 | -10 | -47.6 | -158.68 | -13265.6 | 13266.6344 | 14:48
40 | 331.5 | 0 | -20 | -31.2 | -326.6 | -13229.6 | 13233.6676 | 14:48
41 | 379 | -20 | 0 | 1299.6 | 5.04 | -6415.2 | 6545.5158 | 14:48
42 | 379 | -10 | 0 | 656.5 | 6.22 | -6550.1 | 6582.9204 | 14:48
43 | 379 | 0 | 0 | 12.3 | 8.72 | -6597.6 | 6597.6172 | 14:49
44 | 379 | 10 | 0 | -627.8 | 11.29 | -6560.4 | 6590.38 | 14:49
45 | 379 | 20 | 0 | -1261 | 13.54 | -6437 | 6559.3653 | 14:49
46 | 379 | 0 | 20 | 3.4 | 231.51 | -6454.4 | 6458.5515 | 14:50
47 | 379 | 0 | 10 | 4.5 | 120.74 | -6573.3 | 6574.4103 | 14:50
48 | 379 | 0 | 0 | 9.4 | 9.17 | -6599.2 | 6599.2131 | 14:50
49 | 379 | 0 | -10 | 11.8 | -101.62 | -6542.4 | 6543.1998 | 14:50
50 | 379 | 0 | -20 | 20.8 | -212.98 | -6392.7 | 6396.2807 | 14:51
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
| Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt
3 | 20 | 0 | 0 | -173.5 | -23.44 | -11145.7 | 11147.075
8 | 20 | 0 | 0 | -171.7 | -24.03 | -11144.5 | 11145.8485
13 | 67.5 | 0 | 0 | -240.4 | -44.31 | -17915.5 | 17917.1676
18 | 67.5 | 0 | 0 | -237.9 | -45.29 | -17914 | 17915.6368
23 | 199.5 | 0 | 0 | -186.8 | -49.3 | -19963 | 19963.9348
28 | 199.5 | 0 | 0 | -186.7 | -49.35 | -19963.2 | 19964.134
33 | 331.5 | 0 | 0 | -55.3 | -0.52 | -13309.9 | 13310.0149
38 | 331.5 | 0 | 0 | -57.9 | 0.87 | -13311.3 | 13311.426
43 | 379 | 0 | 0 | 12.3 | 8.72 | -6597.6 | 6597.6172
48 | 379 | 0 | 0 | 9.4 | 9.17 | -6599.2 | 6599.2131

## HallScan@1T
XYZ MAPPING | ABLE UNIFORMI | Y DA | A FILE
est Engineer: Scott Choroba
Date: | ue, Feb 03, 2015
Magnet Model: 14285 | RIUMF
Magnet Serial No:14285
MAGNE | SCAN A | 1
| Z-axis | Vertical | Horizontal | Horizontal |  | Vertical | Z-axis
|  |  |  | Tesla |  | Tesla | Tesla |  | Clock was wrong, started in AM not PM
Step | Z (mm) | Y (mm) | X (mm) | Bx | 10*Bx | By | Bz | Bt | Time
2 | 20 | 0 | 0 | -0.00267 | -0.0267 | -0.04359 | -0.39919 | 0.40157 | 23:56
4 | 0 | 0 | 0 | -0.00266 | -0.0266 | -0.04826 | -0.46667 | 0.46917 | 23:57
4 | 0 | 0 | 0 | -0.00268 | -0.0268 | -0.05325 | -0.54156 | 0.54418 | 23:57
6 | 0 | 0 | 0 | -0.00324 | -0.0324 | -0.05816 | -0.6214 | 0.62412 | 23:57
6 | 0 | 0 | 0 | -0.00393 | -0.0393 | -0.06445 | -0.70181 | 0.70477 | 23:58
8 | 0 | 0 | 0 | -0.00542 | -0.0542 | -0.07009 | -0.77785 | 0.78102 | 23:58
8 | 80 | 0 | 0 | -0.00611 | -0.0611 | -0.07657 | -0.84483 | 0.84832 | 23:58
9 | 90 | 0 | 0 | -0.00663 | -0.0663 | -0.08158 | -0.89918 | 0.9029 | 23:59
10 | 90 | 0 | 0 | -0.00701 | -0.0701 | -0.08167 | -0.89916 | 0.90289 | 23:59
11 | 100 | 0 | 0 | -0.00717 | -0.0717 | -0.08589 | -0.93944 | 0.94339 | 0:00
11 | 0 | 0 | 0 | -0.00718 | -0.0718 | -0.08878 | -0.96642 | 0.97052 | 0:01
13 | 120 | 0 | 0 | -0.00713 | -0.0713 | -0.09048 | -0.98249 | 0.98667 | 0:04
13 | 130 | 0 | 0 | -0.00713 | -0.0713 | -0.09157 | -0.99096 | 0.99521 | 0:06
15 | 140 | 0 | 0 | -0.0071 | -0.071 | -0.09213 | -0.99489 | 0.99917 | 0:07
15 | 150 | 0 | 0 | -0.00703 | -0.0703 | -0.09217 | -0.99653 | 1.0008 | 0:07
16 | 150 | 0 | 0 | -0.00698 | -0.0698 | -0.09232 | -0.99653 | 1.0008 | 0:08
17 | 160 | 0 | 0 | -0.00693 | -0.0693 | -0.09261 | -0.99722 | 1.0015 | 0:08
17 | 160 | 0 | 0 | -0.00693 | -0.0693 | -0.09254 | -0.99723 | 1.0015 | 0:09
18 | 170 | 0 | 0 | -0.00691 | -0.0691 | -0.09291 | -0.99759 | 1.0019 | 0:11
18 | 0 | 0 | 0 | -0.00692 | -0.0692 | -0.09314 | -0.99782 | 1.0022 | 0:11
19 | 190 | 0 | 0 | -0.00688 | -0.0688 | -0.09326 | -0.99797 | 1.0023 | 0:12
22 | 210 | 0 | 0 | -0.00674 | -0.0674 | -0.09331 | -0.99806 | 1.0024 | 0:16
23 | 220 | 0 | 0 | -0.00671 | -0.0671 | -0.09336 | -0.998 | 1.0024 | 0:17
24 | 230 | 0 | 0 | -0.00657 | -0.0657 | -0.09322 | -0.99788 | 1.0022 | 0:19
1 | 0 | 0 | 0 | -0.00185 | -0.0185 | -0.03509 | -0.28994 | 0.29206 | 0:39
2 | 10 | 0 | 0 | -0.00221 | -0.0221 | -0.03897 | -0.34022 | 0.34245 | 0:39
3 | 20 | 0 | 0 | -0.00266 | -0.0266 | -0.04312 | -0.39917 | 0.4015 | 0:40
4 | 30 | 0 | 0 | -0.00323 | -0.0323 | -0.04737 | -0.46662 | 0.46903 | 0:40
5 | 40 | 0 | 0 | -0.0039 | -0.039 | -0.05226 | -0.54153 | 0.54406 | 0:40
6 | 50 | 0 | 0 | -0.00464 | -0.0464 | -0.05708 | -0.62134 | 0.62397 | 0:41
7 | 60 | 0 | 0 | -0.00537 | -0.0537 | -0.06255 | -0.70176 | 0.70456 | 0:41
8 | 70 | 0 | 0 | -0.00606 | -0.0606 | -0.06869 | -0.77785 | 0.7809 | 0:41
9 | 80 | 0 | 0 | -0.00657 | -0.0657 | -0.0753 | -0.84477 | 0.84815 | 0:42
10 | 90 | 0 | 0 | -0.00698 | -0.0698 | -0.08097 | -0.89914 | 0.9028 | 0:42
11 | 100 | 0 | 0 | -0.00715 | -0.0715 | -0.08504 | -0.93942 | 0.94329 | 0:42
12 | 110 | 0 | 0 | -0.00714 | -0.0714 | -0.08768 | -0.96643 | 0.97043 | 0:42
13 | 120 | 0 | 0 | -0.00713 | -0.0713 | -0.08991 | -0.98247 | 0.9866 | 0:43
14 | 130 | 0 | 0 | -0.00709 | -0.0709 | -0.09097 | -0.99093 | 0.99512 | 0:43
15 | 140 | 0 | 0 | -0.00703 | -0.0703 | -0.09113 | -0.99485 | 0.99904 | 0:43
16 | 150 | 0 | 0 | -0.00698 | -0.0698 | -0.09152 | -0.99651 | 1.0007 | 0:44
17 | 160 | 0 | 0 | -0.00695 | -0.0695 | -0.09194 | -0.99721 | 1.0015 | 0:44
18 | 170 | 0 | 0 | -0.00691 | -0.0691 | -0.09233 | -0.99756 | 1.0018 | 0:44
19 | 180 | 0 | 0 | -0.00688 | -0.0688 | -0.09262 | -0.9978 | 1.0021 | 0:45
3 | -195 | 0 | 0 | -0.00209 | -0.0209 | -0.03701 | -0.32182 | 0.32395 | 0:48
4 | -190 | 0 | 0 | -0.00229 | -0.0229 | -0.03927 | -0.34869 | 0.3509 | 0:48
5 | -185 | 0 | 0 | -0.00251 | -0.0251 | -0.04146 | -0.37772 | 0.38 | 0:49
6 | -180 | 0 | 0 | -0.00275 | -0.0275 | -0.04396 | -0.40896 | 0.41132 | 0:49
7 | -175 | 0 | 0 | -0.00303 | -0.0303 | -0.04595 | -0.44231 | 0.4447 | 0:49
8 | -170 | 0 | 0 | -0.00333 | -0.0333 | -0.04844 | -0.47765 | 0.48011 | 0:50
9 | -165 | 0 | 0 | -0.00368 | -0.0368 | -0.05064 | -0.51481 | 0.51731 | 0:50
10 | -160 | 0 | 0 | -0.00396 | -0.0396 | -0.05192 | -0.55342 | 0.55586 | 0:50
11 | -155 | 0 | 0 | -0.00439 | -0.0439 | -0.05498 | -0.59329 | 0.59585 | 0:51
12 | -150 | 0 | 0 | -0.00471 | -0.0471 | -0.05697 | -0.63364 | 0.63621 | 0:51
13 | -145 | 0 | 0 | -0.00513 | -0.0513 | -0.06056 | -0.67403 | 0.67676 | 0:51
14 | -140 | 0 | 0 | -0.00552 | -0.0552 | -0.06369 | -0.71376 | 0.71662 | 0:51
15 | -135 | 0 | 0 | -0.00583 | -0.0583 | -0.06671 | -0.75217 | 0.75514 | 0:52
16 | -130 | 0 | 0 | -0.00616 | -0.0616 | -0.06998 | -0.78875 | 0.79187 | 0:52
17 | -125 | 0 | 0 | -0.00643 | -0.0643 | -0.07294 | -0.8228 | 0.82605 | 0:52
18 | -120 | 0 | 0 | -0.00663 | -0.0663 | -0.07591 | -0.85393 | 0.85732 | 0:53
19 | -115 | 0 | 0 | -0.00682 | -0.0682 | -0.07866 | -0.88176 | 0.88529 | 0:53
20 | -110 | 0 | 0 | -0.00693 | -0.0693 | -0.0806 | -0.90614 | 0.90974 | 0:53
21 | -105 | 0 | 0 | -0.00705 | -0.0705 | -0.08369 | -0.92697 | 0.93077 | 0:53
22 | -100 | 0 | 0 | -0.0071 | -0.071 | -0.08519 | -0.94428 | 0.94814 | 0:54
23 | -95 | 0 | 0 | -0.00714 | -0.0714 | -0.08712 | -0.95835 | 0.96233 | 0:54
24 | -90 | 0 | 0 | -0.00722 | -0.0722 | -0.08805 | -0.96943 | 0.97345 | 0:54
25 | -85 | 0 | 0 | -0.00719 | -0.0719 | -0.08916 | -0.9779 | 0.98198 | 0:55
26 | -80 | 0 | 0 | -0.00717 | -0.0717 | -0.08989 | -0.98417 | 0.98829 | 0:55
27 | -75 | 0 | 0 | -0.00712 | -0.0712 | -0.09058 | -0.98864 | 0.99281 | 0:55
28 | -70 | 0 | 0 | -0.00709 | -0.0709 | -0.09077 | -0.99174 | 0.99591 | 0:56
29 | -65 | 0 | 0 | -0.00708 | -0.0708 | -0.09103 | -0.99382 | 0.99801 | 0:56
30 | -60 | 0 | 0 | -0.00706 | -0.0706 | -0.09109 | -0.9952 | 0.99938 | 0:56
31 | -55 | 0 | 0 | -0.00701 | -0.0701 | -0.09122 | -0.99608 | 1.0003 | 0:56
32 | -50 | 0 | 0 | -0.00701 | -0.0701 | -0.09142 | -0.99664 | 1.0008 | 0:57
33 | -45 | 0 | 0 | -0.00699 | -0.0699 | -0.09172 | -0.99699 | 1.0012 | 0:57
34 | -40 | 0 | 0 | -0.00697 | -0.0697 | -0.0917 | -0.99725 | 1.0015 | 0:57
35 | -35 | 0 | 0 | -0.00696 | -0.0696 | -0.09209 | -0.99744 | 1.0017 | 0:58
36 | -30 | 0 | 0 | -0.00693 | -0.0693 | -0.09212 | -0.99758 | 1.0018 | 0:58
37 | -25 | 0 | 0 | -0.00691 | -0.0691 | -0.0925 | -0.9977 | 1.002 | 0:58
38 | -20 | 0 | 0 | -0.00688 | -0.0688 | -0.09264 | -0.99781 | 1.0021 | 0:58
39 | -15 | 0 | 0 | -0.00687 | -0.0687 | -0.09253 | -0.99789 | 1.0022 | 0:59
40 | -10 | 0 | 0 | -0.00687 | -0.0687 | -0.0927 | -0.99794 | 1.0023 | 0:59
41 | -5 | 0 | 0 | -0.00682 | -0.0682 | -0.09271 | -0.99798 | 1.0023 | 0:59
42 | 0 | 0 | 0 | -0.00682 | -0.0682 | -0.0927 | -0.99801 | 1.0023 | 1:00
43 | 5 | 0 | 0 | -0.0068 | -0.068 | -0.0929 | -0.99803 | 1.0024 | 1:00
44 | 10 | 0 | 0 | -0.00678 | -0.0678 | -0.09291 | -0.99803 | 1.0024 | 1:00
45 | 15 | 0 | 0 | -0.00675 | -0.0675 | -0.09307 | -0.998 | 1.0024 | 1:01
46 | 20 | 0 | 0 | -0.00674 | -0.0674 | -0.09298 | -0.99796 | 1.0023 | 1:01
47 | 25 | 0 | 0 | -0.00671 | -0.0671 | -0.0929 | -0.99791 | 1.0022 | 1:01
48 | 30 | 0 | 0 | -0.00669 | -0.0669 | -0.09284 | -0.99783 | 1.0022 | 1:01
49 | 35 | 0 | 0 | -0.00668 | -0.0668 | -0.09257 | -0.99773 | 1.002 | 1:02
50 | 40 | 0 | 0 | -0.00665 | -0.0665 | -0.09251 | -0.99764 | 1.0019 | 1:02
51 | 45 | 0 | 0 | -0.00663 | -0.0663 | -0.09237 | -0.99752 | 1.0018 | 1:02
52 | 50 | 0 | 0 | -0.00662 | -0.0662 | -0.0923 | -0.99733 | 1.0016 | 1:03
53 | 55 | 0 | 0 | -0.00662 | -0.0662 | -0.0923 | -0.99705 | 1.0013 | 1:03
54 | 60 | 0 | 0 | -0.00659 | -0.0659 | -0.09212 | -0.99659 | 1.0009 | 1:03
55 | 65 | 0 | 0 | -0.00659 | -0.0659 | -0.09175 | -0.99586 | 1.0001 | 1:04
56 | 70 | 0 | 0 | -0.00654 | -0.0654 | -0.09159 | -0.99468 | 0.99891 | 1:04
57 | 75 | 0 | 0 | -0.00652 | -0.0652 | -0.09119 | -0.99283 | 0.99703 | 1:04
58 | 80 | 0 | 0 | -0.00651 | -0.0651 | -0.09089 | -0.99003 | 0.99421 | 1:04
59 | 85 | 0 | 0 | -0.00647 | -0.0647 | -0.09018 | -0.98594 | 0.99008 | 1:05
60 | 90 | 0 | 0 | -0.00641 | -0.0641 | -0.08914 | -0.98017 | 0.98424 | 1:05
61 | 95 | 0 | 0 | -0.00635 | -0.0635 | -0.08809 | -0.97231 | 0.97631 | 1:05
62 | 100 | 0 | 0 | -0.00626 | -0.0626 | -0.08668 | -0.96191 | 0.96583 | 1:06
63 | 105 | 0 | 0 | -0.00615 | -0.0615 | -0.08486 | -0.94859 | 0.9524 | 1:06
64 | 110 | 0 | 0 | -0.00602 | -0.0602 | -0.08261 | -0.93207 | 0.93574 | 1:06
65 | 115 | 0 | 0 | -0.00589 | -0.0589 | -0.07958 | -0.91211 | 0.91559 | 1:06
66 | 120 | 0 | 0 | -0.00568 | -0.0568 | -0.07609 | -0.88856 | 0.89183 | 1:07
67 | 125 | 0 | 0 | -0.00546 | -0.0546 | -0.07285 | -0.86151 | 0.8646 | 1:07
68 | 130 | 0 | 0 | -0.00522 | -0.0522 | -0.06939 | -0.83117 | 0.83408 | 1:07
69 | 135 | 0 | 0 | -0.00493 | -0.0493 | -0.06536 | -0.79769 | 0.80038 | 1:08
70 | 140 | 0 | 0 | -0.00466 | -0.0466 | -0.06105 | -0.76169 | 0.76415 | 1:08
71 | 150 | 0 | 0 | -0.00393 | -0.0393 | -0.05239 | -0.6841 | 0.68611 | 1:08
72 | 155 | 0 | 0 | -0.00357 | -0.0357 | -0.04809 | -0.64376 | 0.64556 | 1:09
73 | 160 | 0 | 0 | -0.00314 | -0.0314 | -0.04333 | -0.60323 | 0.60479 | 1:09
74 | 165 | 0 | 0 | -0.00281 | -0.0281 | -0.04046 | -0.56326 | 0.56472 | 1:09
75 | 170 | 0 | 0 | -0.00246 | -0.0246 | -0.03712 | -0.5242 | 0.52552 | 1:09
76 | 175 | 0 | 0 | -0.00209 | -0.0209 | -0.03429 | -0.4866 | 0.48781 | 1:10
77 | 180 | 0 | 0 | -0.00185 | -0.0185 | -0.03223 | -0.45079 | 0.45194 | 1:10
78 | 185 | 0 | 0 | -0.00149 | -0.0149 | -0.02994 | -0.41685 | 0.41793 | 1:10
79 | 190 | 0 | 0 | -0.00126 | -0.0126 | -0.02846 | -0.38509 | 0.38614 | 1:11
80 | 195 | 0 | 0 | -0.00102 | -0.0102 | -0.02668 | -0.35548 | 0.35648 | 1:11
81 | 200 | 0 | 0 | -0.00083 | -0.0083 | -0.02487 | -0.32803 | 0.32897 | 1:11
82 | 205 | 0 | 0 | -0.00069 | -0.0069 | -0.02359 | -0.30276 | 0.30368 | 1:12
83 | 210 | 0 | 0 | -0.00055 | -0.0055 | -0.02239 | -0.27949 | 0.28039 | 1:12

## Zmod.No.Back
XYZ MAPPIN | TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Thu, Feb 05, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
5A INTO Z MOD COIL. NO CURRENT ON OTHER MOD COILS OR ON Z COIL.
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
Step | Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt | Time
1 | 1 | 0 | 0 | 1.38 | 4.29 | -7.28 | 8.5619 | 14:52
2 | 10 | 0 | 0 | 1.4 | 4.24 | -4.01 | 6.0015 | 14:52
3 | 20 | 0 | 0 | 1.54 | 4.54 | 1.42 | 5 | 14:53
4 | 30 | 0 | 0 | 1.7 | 4.76 | 8.35 | 9.7606 | 14:53
5 | 40 | 0 | 0 | 1.8 | 5.09 | 16.25 | 17.1234 | 14:53
6 | 50 | 0 | 0 | 1.78 | 5.39 | 23.65 | 24.3217 | 14:54
7 | 60 | 0 | 0 | 1.6 | 5.52 | 28.8 | 29.3678 | 14:54
8 | 70 | 0 | 0 | 1.24 | 5.47 | 30.11 | 30.6279 | 14:54
9 | 80 | 0 | 0 | 0.76 | 5.17 | 26.76 | 27.2654 | 14:54
10 | 90 | 0 | 0 | 0.3 | 4.73 | 18.62 | 19.2137 | 14:55
11 | 100 | 0 | 0 | -0.06 | 4.06 | 6.64 | 7.7831 | 14:55
12 | 110 | 0 | 0 | -0.28 | 3.4 | -7.59 | 8.3214 | 14:55
13 | 120 | 0 | 0 | -0.3 | 2.78 | -22.23 | 22.4052 | 14:56
14 | 130 | 0 | 0 | -0.13 | 2.32 | -35.35 | 35.4263 | 14:56
15 | 140 | 0 | 0 | 0.17 | 1.99 | -46.1 | 46.1432 | 14:56
16 | 150 | 0 | 0 | 0.53 | 1.83 | -53.96 | 53.9936 | 14:56
17 | 160 | 0 | 0 | 0.83 | 1.7 | -59.42 | 59.4501 | 14:57
18 | 170 | 0 | 0 | 1.07 | 1.75 | -63.07 | 63.1033 | 14:57
19 | 180 | 0 | 0 | 1.27 | 1.7 | -65.15 | 65.1845 | 14:57
20 | 190 | 0 | 0 | 1.41 | 1.83 | -66.44 | 66.4802 | 14:57
21 | 200 | 0 | 0 | 1.5 | 1.93 | -66.76 | 66.8047 | 14:58
22 | 210 | 0 | 0 | 1.56 | 1.93 | -66.51 | 66.5563 | 14:58
23 | 220 | 0 | 0 | 1.62 | 2 | -65.41 | 65.4606 | 14:58
24 | 230 | 0 | 0 | 1.68 | 2.24 | -63.48 | 63.5417 | 14:59
25 | 240 | 0 | 0 | 1.75 | 2.49 | -60.08 | 60.157 | 14:59
26 | 250 | 0 | 0 | 1.84 | 2.81 | -54.8 | 54.9028 | 14:59
27 | 260 | 0 | 0 | 1.97 | 3.18 | -47.17 | 47.3181 | 14:59
28 | 270 | 0 | 0 | 2.12 | 3.91 | -36.79 | 37.0579 | 15:00
29 | 280 | 0 | 0 | 2.25 | 4.61 | -23.73 | 24.2781 | 15:00
30 | 290 | 0 | 0 | 2.31 | 5.39 | -9.04 | 10.7754 | 15:00
31 | 300 | 0 | 0 | 2.27 | 6.04 | 5.55 | 8.511 | 15:01
32 | 310 | 0 | 0 | 2.08 | 6.51 | 18.12 | 19.366 | 15:01
33 | 320 | 0 | 0 | 1.78 | 6.71 | 26.88 | 27.762 | 15:01
34 | 330 | 0 | 0 | 1.39 | 6.41 | 31.07 | 31.7548 | 15:01
35 | 340 | 0 | 0 | 1.02 | 5.93 | 30.52 | 31.1075 | 15:02
36 | 350 | 0 | 0 | 0.68 | 5.24 | 25.69 | 26.2278 | 15:02
37 | 360 | 0 | 0 | 0.47 | 4.56 | 18.29 | 18.8557 | 15:02
38 | 370 | 0 | 0 | 0.44 | 4.01 | 10.18 | 10.9502 | 15:02
39 | 380 | 0 | 0 | 0.5 | 3.68 | 2.83 | 4.6692 | 15:03
40 | 390 | 0 | 0 | 0.65 | 3.69 | -2.97 | 4.7812 | 15:03
41 | 400 | 0 | 0 | 0.8 | 3.71 | -6.98 | 7.9451 | 15:03
42 | 410 | 0 | 0 | 0.97 | 3.69 | -9.65 | 10.3769 | 15:04
43 | 0 | 0 | 0 | 2.15 | 4.11 | -61.99 | 62.1633 | 15:04
43 | 0 | 0 | 0 | 1.63 | 4.02 | -10.08 | 10.9738 | 15:04

## Ymod.NO.Back
XYZ MAPPIN | TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date:Wed, Feb 04, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
5A INTO Y MOD COIL. NO CURRENT ON OTHER MOD COILS OR ON Z COIL.
MOVIN | PROBE DOWN TO UP AS SEEN FROM BORE OPPOSITE OF RED RELIEFE VALVE
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
Step | Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt | Time
0 | 207.5 | -13 | 0 | -9.52 | -2.66 | -49.18 | 50.1635 | 16:13
0 | 207.5 | -12 | 0 | -9.5 | -2.75 | -49.18 | 50.1646 | 16:13
0 | 207.5 | -11 | 0 | -9.47 | -2.65 | -49.19 | 50.1633 | 16:13
0 | 207.5 | -10 | 0 | -9.45 | -2.65 | -49.21 | 50.1792 | 16:14
0 | 207.5 | -9 | 0 | -9.4 | -2.71 | -49.22 | 50.1828 | 16:14
0 | 207.5 | -8 | 0 | -9.38 | -2.65 | -49.21 | 50.166 | 16:14
0 | 207.5 | -7 | 0 | -9.36 | -2.58 | -49.21 | 50.1586 | 16:14
0 | 207.5 | -6 | 0 | -9.35 | -2.68 | -49.23 | 50.1816 | 16:14
0 | 207.5 | -5 | 0 | -9.32 | -2.64 | -49.22 | 50.1641 | 16:14
0 | 207.5 | -4 | 0 | -9.3 | -2.58 | -49.22 | 50.1573 | 16:15
0 | 207.5 | -3 | 0 | -9.29 | -2.69 | -49.23 | 50.171 | 16:15
0 | 207.5 | -2 | 0 | -9.28 | -2.7 | -49.21 | 50.1501 | 16:15
0 | 207.5 | -1 | 0 | -9.27 | -2.72 | -49.21 | 50.1493 | 16:15
0 | 207.5 | 0 | 0 | -9.26 | -2.69 | -49.2 | 50.1361 | 16:15
0 | 207.5 | 1 | 0 | -9.27 | -2.72 | -49.2 | 50.1395 | 16:16
0 | 207.5 | 2 | 0 | -9.29 | -2.69 | -49.19 | 50.1318 | 16:16
0 | 207.5 | 3 | 0 | -9.28 | -2.73 | -49.16 | 50.1027 | 16:16
0 | 207.5 | 4 | 0 | -9.3 | -2.72 | -49.16 | 50.1058 | 16:16
0 | 207.5 | 5 | 0 | -9.3 | -2.68 | -49.15 | 50.0939 | 16:16
0 | 207.5 | 6 | 0 | -9.31 | -2.66 | -49.13 | 50.075 | 16:17
0 | 207.5 | 7 | 0 | -9.34 | -2.65 | -49.1 | 50.0507 | 16:17
0 | 207.5 | 8 | 0 | -9.33 | -2.65 | -49.08 | 50.0292 | 16:17
0 | 207.5 | 9 | 0 | -9.42 | -2.65 | -49.07 | 50.0362 | 16:17
0 | 207.5 | 8 | 0 | -9.41 | -2.7 | -49.1 | 50.0664 | 16:18
0 | 207.5 | 10 | 0 | -9.47 | -2.7 | -49.06 | 50.0385 | 16:18
0 | 207.5 | 11 | 0 | -9.48 | -2.69 | -49.04 | 50.0203 | 16:18
0 | 207.5 | 12 | 0 | -9.51 | -2.72 | -49.03 | 50.0178 | 16:18
0 | 207.5 | 13 | 0 | -9.54 | -2.71 | -49 | 49.9936 | 16:18
|  | MOVIN | PROBE LEFT TO RI | HT AS SEEN FROM BORE OPPOSITE OF RED RELIEFE VALVE
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
Step | Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt | Time
0 | 207.5 | 0 | -11.5 | -9.27 | -2.61 | -49.17 | 50.1042 | 16:19
0 | 207.5 | 0 | -10.5 | -9.27 | -2.61 | -49.19 | 50.1239 | 16:19
0 | 207.5 | 0 | -9.5 | -9.27 | -2.59 | -49.18 | 50.113 | 16:20
0 | 207.5 | 0 | -8.5 | -9.27 | -2.56 | -49.19 | 50.1213 | 16:20
0 | 207.5 | 0 | -7.5 | -9.29 | -2.58 | -49.19 | 50.126 | 16:20
0 | 207.5 | 0 | -6.5 | -9.3 | -2.65 | -49.21 | 50.1511 | 16:20
0 | 207.5 | 0 | -5.5 | -9.31 | -2.58 | -49.21 | 50.1493 | 16:20
0 | 207.5 | 0 | -4.5 | -9.31 | -2.7 | -49.21 | 50.1557 | 16:20
0 | 207.5 | 0 | -3.5 | -9.31 | -2.74 | -49.21 | 50.1578 | 16:21
0 | 207.5 | 0 | -2.5 | -9.3 | -2.76 | -49.22 | 50.1669 | 16:21
0 | 207.5 | 0 | -1.5 | -9.3 | -2.75 | -49.21 | 50.1565 | 16:21
0 | 207.5 | 0 | -0.5 | -9.31 | -2.69 | -49.2 | 50.1453 | 16:22
0 | 207.5 | 0 | 0.5 | -9.31 | -2.66 | -49.2 | 50.1437 | 16:22
0 | 207.5 | 0 | 1.5 | -9.29 | -2.66 | -49.19 | 50.1302 | 16:22
0 | 207.5 | 0 | 2.5 | -9.31 | -2.69 | -49.19 | 50.1355 | 16:22
0 | 207.5 | 0 | 3.5 | -9.29 | -2.84 | -49.19 | 50.1401 | 16:22
0 | 207.5 | 0 | 4.5 | -9.27 | -2.75 | -49.18 | 50.1215 | 16:23
0 | 207.5 | 0 | 5.5 | -9.24 | -2.79 | -49.17 | 50.1084 | 16:23
0 | 207.5 | 0 | 6.5 | -9.23 | -2.84 | -49.15 | 50.0897 | 16:23
0 | 207.5 | 0 | 7.5 | -9.22 | -2.77 | -49.13 | 50.0643 | 16:23
0 | 207.5 | 0 | 8.5 | -9.2 | -2.78 | -49.12 | 50.0514 | 16:23
0 | 207.5 | 0 | 9.5 | -9.19 | -2.79 | -49.1 | 50.0305 | 16:24
0 | 207.5 | 0 | 10.5 | -9.18 | -2.89 | -49.08 | 50.0147 | 16:24
0 | 207.5 | 0 | 11.5 | -9.17 | -2.83 | -49.07 | 49.9996 | 16:24
0 | 207.5 | 0 | 12.5 | -9.14 | -2.84 | -49.04 | 49.9653 | 16:24

## Xmod.NO.Back
XYZ MAPPIN | TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Wed, Feb 04, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
5A INTO X MOD COIL. NO CURRENT ON OTHER MOD COILS OR ON Z COIL.
MOVIN | PROBE HORIZONTALLY AS VIEWED FROM THE BORE RI | HT TO LEFT SIDE OPPOSITEOF RED RELIEF VALVE
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
Step | Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt | Time
0 | 207.5 | -1 | 30 | 1.32 | -25.87 | -47.67 | 54.2534 | 15:18
0 | 207.5 | -1 | 29 | 1.31 | -25.66 | -47.72 | 54.1973 | 15:18
0 | 207.5 | -1 | 28 | 1.32 | -25.44 | -47.79 | 54.1555 | 15:19
0 | 207.5 | -1 | 27 | 1.32 | -25.19 | -47.85 | 54.0916 | 15:19
0 | 207.5 | -1 | 26 | 1.32 | -25.02 | -47.91 | 54.0658 | 15:19
0 | 207.5 | -1 | 25 | 1.32 | -24.83 | -47.95 | 54.0136 | 15:19
0 | 207.5 | -1 | 24 | 1.31 | -24.57 | -48 | 53.9389 | 15:20
0 | 207.5 | -1 | 23 | 1.33 | -24.39 | -48.06 | 53.9111 | 15:20
0 | 207.5 | -1 | 22 | 1.32 | -24.21 | -48.09 | 53.8564 | 15:20
0 | 207.5 | -1 | 21 | 1.32 | -24.02 | -48.15 | 53.825 | 15:20
0 | 207.5 | -1 | 20 | 1.3 | -23.83 | -48.17 | 53.7579 | 15:21
0 | 207.5 | -1 | 19 | 1.3 | -23.72 | -48.2 | 53.7361 | 15:21
0 | 207.5 | -1 | 18 | 1.32 | -23.53 | -48.23 | 53.6799 | 15:21
0 | 207.5 | -1 | 17 | 1.32 | -23.42 | -48.26 | 53.6588 | 15:21
0 | 207.5 | -1 | 16 | 1.31 | -23.23 | -48.28 | 53.5939 | 15:22
0 | 207.5 | -1 | 15 | 1.3 | -23.13 | -48.29 | 53.5594 | 15:22
0 | 207.5 | -1 | 14 | 1.3 | -22.99 | -48.32 | 53.5262 | 15:23
0 | 207.5 | -1 | 13 | 1.31 | -22.92 | -48.33 | 53.5054 | 15:23
0 | 207.5 | -1 | 12 | 1.31 | -22.76 | -48.35 | 53.4552 | 15:23
0 | 207.5 | -1 | 11 | 1.29 | -22.68 | -48.36 | 53.4297 | 15:23
0 | 207.5 | -1 | 10 | 1.3 | -22.54 | -48.38 | 53.3888 | 15:24
0 | 207.5 | -1 | 9 | 1.3 | -22.35 | -48.38 | 53.3089 | 15:24
0 | 207.5 | -1 | 8 | 1.29 | -22.39 | -48.39 | 53.3345 | 15:24
0 | 207.5 | -1 | 7 | 1.29 | -22.34 | -48.4 | 53.3226 | 15:24
0 | 207.5 | -1 | 6 | 1.29 | -22.33 | -48.39 | 53.3093 | 15:25
0 | 207.5 | -1 | 5 | 1.3 | -22.23 | -48.39 | 53.2678 | 15:25
0 | 207.5 | -1 | 4 | 1.28 | -22.18 | -48.38 | 53.2373 | 15:25
0 | 207.5 | -1 | 5 | 1.28 | -22.23 | -48.39 | 53.2673 | 15:26
0 | 207.5 | -1 | 4 | 1.3 | -22.21 | -48.39 | 53.2594 | 15:26
0 | 207.5 | -1 | 3 | 1.29 | -22.18 | -48.38 | 53.2376 | 15:26
0 | 207.5 | -1 | 2 | 1.28 | -22.13 | -48.39 | 53.2256 | 15:26
0 | 207.5 | -1 | 1 | 1.29 | -22.06 | -48.38 | 53.1877 | 15:26
0 | 207.5 | -1 | 0 | 1.28 | -22.09 | -48.36 | 53.1817 | 15:27
0 | 207.5 | -1 | -1 | 1.28 | -21.96 | -48.35 | 53.1188 | 15:27
0 | 207.5 | -1 | -2 | 1.26 | -22 | -48.36 | 53.1439 | 15:27
0 | 207.5 | -1 | -3 | 1.28 | -22.04 | -48.35 | 53.1519 | 15:27
0 | 207.5 | -1 | -4 | 1.27 | -22.02 | -48.33 | 53.1252 | 15:28
0 | 207.5 | -1 | -5 | 1.27 | -22.04 | -48.32 | 53.1244 | 15:28
0 | 207.5 | -1 | -6 | 1.26 | -22.05 | -48.3 | 53.1101 | 15:28
0 | 207.5 | -1 | -7 | 1.26 | -22.07 | -48.28 | 53.1002 | 15:29
0 | 207.5 | -1 | -8 | 1.25 | -22.04 | -48.26 | 53.0693 | 15:29
0 | 207.5 | -1 | -9 | 1.26 | -22.11 | -48.24 | 53.0805 | 15:29
0 | 207.5 | -1 | -10 | 1.25 | -22.15 | -48.23 | 53.0878 | 15:29
0 | 207.5 | -1 | -11 | 1.25 | -22.16 | -48.21 | 53.0738 | 15:30
0 | 207.5 | -1 | -12 | 1.24 | -22.2 | -48.18 | 53.0631 | 15:30
0 | 207.5 | -1 | -13 | 1.25 | -22.23 | -48.17 | 53.0668 | 15:30
0 | 207.5 | -1 | -14 | 1.25 | -22.26 | -48.18 | 53.0884 | 15:30
MOVIN | PROBE VERTIALLY AS VIEWED FROM THE BORE DOWN TO UP SIDE OPPOSITEOF RED RELIEF VALVE
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
Step | Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt | Time
0 | 207.5 | -12 | 0 | 0.68 | -26.87 | -48.51 | 55.4588 | 15:54
0 | 207.5 | -11 | 0 | 0.68 | -26.87 | -48.51 | 55.4588 | 15:55
0 | 207.5 | -10 | 0 | 0.67 | -26.88 | -48.52 | 55.4723 | 15:55
0 | 207.5 | -9 | 0 | 0.66 | -26.96 | -48.56 | 55.5459 | 15:55
0 | 207.5 | -8 | 0 | 0.64 | -27.03 | -48.55 | 55.571 | 15:55
0 | 207.5 | -7 | 0 | 0.62 | -27.05 | -48.57 | 55.5979 | 15:56
0 | 207.5 | -6 | 0 | 0.62 | -27.09 | -48.58 | 55.6262 | 15:56
0 | 207.5 | -5 | 0 | 0.59 | -27.07 | -48.59 | 55.6248 | 15:56
0 | 207.5 | -4 | 0 | 0.6 | -27.1 | -48.6 | 55.6483 | 15:56
0 | 207.5 | -3 | 0 | 0.6 | -27.17 | -48.61 | 55.6911 | 15:56
0 | 207.5 | -2 | 0 | 0.6 | -27.12 | -48.62 | 55.6755 | 15:57
0 | 207.5 | -1 | 0 | 0.59 | -27.11 | -48.62 | 55.6705 | 15:57
0 | 207.5 | 0 | 0 | 0.56 | -27.14 | -48.62 | 55.6848 | 15:57
0 | 207.5 | 1 | 0 | 0.57 | -27.11 | -48.63 | 55.679 | 15:57
0 | 207.5 | 2 | 0 | 0.56 | -27.2 | -48.62 | 55.7141 | 15:58
0 | 207.5 | 3 | 0 | 0.55 | -27.15 | -48.63 | 55.6983 | 15:58
0 | 207.5 | 4 | 0 | 0.54 | -27.09 | -48.63 | 55.669 | 15:58
0 | 207.5 | 5 | 0 | 0.53 | -27.17 | -48.62 | 55.6991 | 15:58
0 | 207.5 | 6 | 0 | 0.52 | -27.18 | -48.63 | 55.7127 | 15:59
0 | 207.5 | 7 | 0 | 0.51 | -27.15 | -48.62 | 55.6892 | 15:59
0 | 207.5 | 8 | 0 | 0.52 | -27.05 | -48.61 | 55.6319 | 15:59
0 | 207.5 | 9 | 0 | 0.51 | -27.1 | -48.61 | 55.6561 | 15:59
0 | 207.5 | 10 | 0 | 0.5 | -27.1 | -48.6 | 55.6473 | 15:59
0 | 207.5 | 11 | 0 | 0.47 | -27.01 | -48.58 | 55.5858 | 16:00
0 | 207.5 | 12 | 0 | 0.46 | -26.99 | -48.57 | 55.5672 | 16:00

## YMODAXIALLY
XYZ MAPPIN | TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date:Wed, Feb 04, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
5A INTOZ YMOD, SCANNED LEN | THWISE OF BORE. (AXIALLY) NO CURRENT ON OTHER MOD COILS OR ON Z COIL.
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
Step | Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt | Time
1 | 1 | 0 | 0 | 0.62 | -8.372 | -10.02 | 13.0719 | 19:02
2 | 10 | 0 | 0 | 0.67 | -8.39 | -8.33 | 11.8419 | 19:03
3 | 20 | 0 | 0 | 0.63 | -8.461 | -4.98 | 9.838 | 19:03
4 | 30 | 0 | 0 | 0.55 | -8.523 | -0.02 | 8.5408 | 19:03
5 | 40 | 0 | 0 | 0.4 | -8.561 | 6.64 | 10.8416 | 19:04
6 | 50 | 0 | 0 | 0.2 | -8.579 | 14.36 | 16.7287 | 19:04
7 | 60 | 0 | 0 | -0.07 | -8.534 | 22.47 | 24.0361 | 19:04
8 | 70 | 0 | 0 | -0.44 | -8.384 | 28.87 | 30.066 | 19:04
9 | 80 | 0 | 0 | -0.87 | -8.156 | 31.88 | 32.9183 | 19:05
10 | 100 | 0 | 0 | -2.08 | -7.702 | 25.51 | 26.7284 | 19:05
11 | 110 | 0 | 0 | -2.66 | -7.564 | 16.59 | 18.426 | 19:05
12 | 120 | 0 | 0 | -3.39 | -7.524 | 5.38 | 9.8512 | 19:06
13 | 130 | 0 | 0 | -4.1 | -7.58 | -6.22 | 10.628 | 19:06
14 | 140 | 0 | 0 | -4.9 | -7.732 | -16.41 | 18.7905 | 19:06
15 | 150 | 0 | 0 | -5.69 | -7.933 | -24.07 | 25.9745 | 19:07
16 | 160 | 0 | 0 | -6.49 | -8.105 | -28.98 | 30.7839 | 19:07
17 | 170 | 0 | 0 | -7.24 | -8.263 | -31.56 | 33.4175 | 19:07
18 | 180 | 0 | 0 | -7.94 | -8.344 | -32.38 | 34.3676 | 19:08
19 | 190 | 0 | 0 | -8.49 | -8.389 | -32.49 | 34.6129 | 19:08
20 | 200 | 0 | 0 | -8.92 | -8.391 | -32.23 | 34.4782 | 19:08
21 | 210 | 0 | 0 | -9.23 | -8.373 | -32.03 | 34.3689 | 19:08
22 | 220 | 0 | 0 | -9.38 | -8.361 | -32.06 | 34.4345 | 19:09
23 | 230 | 0 | 0 | -9.36 | -8.363 | -32.33 | 34.6811 | 19:09
24 | 240 | 0 | 0 | -9.22 | -8.359 | -32.61 | 34.9041 | 19:09
25 | 250 | 0 | 0 | -8.89 | -8.386 | -32.46 | 34.6844 | 19:10
26 | 260 | 0 | 0 | -8.44 | -8.449 | -31.09 | 33.3048 | 19:10
27 | 270 | 0 | 0 | -7.87 | -8.524 | -27.81 | 30.1329 | 19:10
28 | 280 | 0 | 0 | -7.23 | -8.631 | -21.9 | 24.6247 | 19:11
29 | 290 | 0 | 0 | -6.56 | -8.705 | -13.18 | 17.1033 | 19:11
30 | 300 | 0 | 0 | -5.84 | -8.749 | -2.17 | 10.7406 | 19:11
31 | 310 | 0 | 0 | -5.06 | -8.682 | 9.73 | 13.9876 | 19:12
32 | 320 | 0 | 0 | -4.29 | -8.522 | 20.48 | 22.5933 | 19:12
33 | 330 | 0 | 0 | -3.49 | -8.298 | 28.56 | 29.9451 | 19:12
34 | 340 | 0 | 0 | -2.71 | -8.062 | 32.55 | 33.6429 | 19:12
35 | 350 | 0 | 0 | -1.95 | -7.836 | 32.27 | 33.265 | 19:13
36 | 360 | 0 | 0 | -1.23 | -7.688 | 27.81 | 28.8793 | 19:13
37 | 370 | 0 | 0 | -0.66 | -7.639 | 20.51 | 21.8963 | 19:13
38 | 380 | 0 | 0 | -0.17 | -7.687 | 12.23 | 14.4462 | 19:14
39 | 390 | 0 | 0 | 0.2 | -7.789 | 4.58 | 9.038 | 19:14
40 | 400 | 0 | 0 | 0.46 | -7.937 | -1.7 | 8.13 | 19:14
41 | 0 | 0 | 0 | -8.5 | -8.404 | -31.99 | 34.1502 | 19:15
41 | 0 | 0 | 0 | 0.33 | -8.369 | -9.9 | 12.9676 | 19:15

## XMODAXIALLY
XYZ MAPPIN | TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Wed, Feb 04, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
5A INTO X MOD COIL, SCANNED LEN | THWISE OF BORE. (AXIALLY) NO CURRENT ON OTHER MOD COILS OR ON Z COIL.
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
Step | Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt | Time
1 | 1 | 0 | 0 | 0.4 | -8.073 | -14.74 | 16.8107 | 17:43
2 | 10 | 0 | 0 | 0.4 | -8.136 | -13.63 | 15.8786 | 17:43
3 | 20 | 0 | 0 | 0.41 | -8.219 | -11.32 | 13.9951 | 17:44
4 | 30 | 0 | 0 | 0.41 | -8.332 | -7.52 | 11.2312 | 17:44
5 | 40 | 0 | 0 | 0.42 | -8.482 | -2.31 | 8.801 | 17:44
6 | 50 | 0 | 0 | 0.39 | -8.598 | 4.12 | 9.5421 | 17:45
7 | 60 | 0 | 0 | 0.33 | -8.705 | 10.6 | 13.7203 | 17:45
8 | 70 | 0 | 0 | 0.25 | -8.764 | 15.47 | 17.7818 | 17:45
9 | 80 | 0 | 0 | 0.21 | -8.803 | 17.28 | 19.3942 | 17:46
10 | 100 | 0 | 0 | 0.19 | -9.065 | 9.03 | 12.7965 | 17:46
11 | 110 | 0 | 0 | 0.23 | -9.358 | -0.46 | 9.3721 | 17:46
12 | 120 | 0 | 0 | 0.31 | -9.771 | -11.94 | 15.4315 | 17:47
13 | 130 | 0 | 0 | 0.41 | -10.318 | -23.5 | 25.6686 | 17:47
14 | 140 | 0 | 0 | 0.47 | -10.962 | -33.7 | 35.4412 | 17:47
15 | 150 | 0 | 0 | 0.55 | -11.618 | -41.31 | 42.9162 | 17:47
16 | 160 | 0 | 0 | 0.58 | -12.236 | -46.07 | 47.6708 | 17:48
17 | 170 | 0 | 0 | 0.57 | -12.782 | -48.58 | 50.2366 | 17:48
18 | 180 | 0 | 0 | 0.54 | -13.206 | -49.35 | 51.0893 | 17:48
19 | 190 | 0 | 0 | 0.52 | -13.506 | -49.41 | 51.2253 | 17:49
20 | 200 | 0 | 0 | 0.51 | -13.712 | -49.12 | 51.0005 | 17:49
21 | 210 | 0 | 0 | 0.5 | -13.819 | -48.88 | 50.7983 | 17:49
22 | 220 | 0 | 0 | 0.49 | -13.832 | -48.9 | 50.821 | 17:50
23 | 230 | 0 | 0 | 0.49 | -13.773 | -49.17 | 51.0649 | 17:50
24 | 240 | 0 | 0 | 0.49 | -13.632 | -49.45 | 51.2969 | 17:50
25 | 250 | 0 | 0 | 0.5 | -13.459 | -49.3 | 51.1066 | 17:51
26 | 260 | 0 | 0 | 0.48 | -13.208 | -47.95 | 49.7382 | 17:51
27 | 270 | 0 | 0 | 0.46 | -12.942 | -44.71 | 46.5477 | 17:51
28 | 280 | 0 | 0 | 0.39 | -12.612 | -38.92 | 40.9143 | 17:51
29 | 290 | 0 | 0 | 0.31 | -12.221 | -30.25 | 32.6269 | 17:52
30 | 300 | 0 | 0 | 0.19 | -11.776 | -19.26 | 22.5756 | 17:52
31 | 310 | 0 | 0 | 0.08 | -11.247 | -7.3 | 13.4086 | 17:53
32 | 320 | 0 | 0 | 0.02 | -10.64 | 3.79 | 11.2949 | 17:53
33 | 330 | 0 | 0 | -0.02 | -9.985 | 12.41 | 15.9282 | 17:53
34 | 340 | 0 | 0 | 0.02 | -9.364 | 17.3 | 19.6717 | 17:54
35 | 350 | 0 | 0 | 0.1 | -8.845 | 18.09 | 20.1368 | 17:54
36 | 360 | 0 | 0 | 0.23 | -8.425 | 15.11 | 17.3016 | 17:54
37 | 370 | 0 | 0 | 0.38 | -8.136 | 9.28 | 12.3474 | 17:54
38 | 380 | 0 | 0 | 0.5 | -8.004 | 2.51 | 8.4032 | 17:55
39 | 390 | 0 | 0 | 0.59 | -7.967 | -3.74 | 8.8209 | 17:55
40 | 400 | 0 | 0 | 0.62 | -7.984 | -8.76 | 11.8687 | 17:55
41 | 0 | 0 | 0 | 0.38 | -12.895 | -49.18 | 50.8439 | 17:56

## ZGRADIENTII
XYZ MAPPING TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Wed, Feb 04, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
5A INTOZ GRAD COIL, SCANNED LENGTHWISE OF BORE. (AXIALLY) NO CURRENT ON OTHER MOD COILS OR ON Z COIL.
| Z-axis | Vertical | Horizontal | Horizontal | Vertical | Z-axis
|  |  |  | Gauss | Gauss | Gauss
Step | Z (mm) | Y (mm) | X (mm) | Bx | By | Bz | Bt | Time
1 | 1 | 0 | 0 | 0.4 | -8.311 | -8.09 | 11.6052 | 18:46
2 | 10 | 0 | 0 | 0.41 | -8.353 | -5.93 | 10.2521 | 18:46
3 | 20 | 0 | 0 | 0.43 | -8.421 | -2.09 | 8.6871 | 18:47
4 | 30 | 0 | 0 | 0.43 | -8.482 | 3.29 | 9.1079 | 18:47
5 | 40 | 0 | 0 | 0.41 | -8.542 | 10.56 | 13.5885 | 18:47
6 | 50 | 0 | 0 | 0.38 | -8.561 | 19.09 | 20.9252 | 18:48
7 | 60 | 0 | 0 | 0.32 | -8.524 | 27.72 | 29.0027 | 18:48
8 | 70 | 0 | 0 | 0.23 | -8.371 | 34.85 | 35.842 | 18:48
9 | 80 | 0 | 0 | 0.16 | -8.147 | 38.56 | 39.4116 | 18:48
10 | 100 | 0 | 0 | 0.13 | -7.667 | 33.04 | 33.9182 | 18:49
11 | 110 | 0 | 0 | 0.16 | -7.525 | 24.18 | 25.3244 | 18:49
12 | 120 | 0 | 0 | 0.22 | -7.481 | 12.86 | 14.8793 | 18:49
13 | 130 | 0 | 0 | 0.34 | -7.539 | 0.92 | 7.6025 | 18:50
14 | 140 | 0 | 0 | 0.44 | -7.692 | -9.8 | 12.466 | 18:50
15 | 150 | 0 | 0 | 0.55 | -7.906 | -18.23 | 19.8781 | 18:50
16 | 160 | 0 | 0 | 0.6 | -8.094 | -23.89 | 25.231 | 18:51
17 | 170 | 0 | 0 | 0.63 | -8.243 | -27.35 | 28.5721 | 18:51
18 | 180 | 0 | 0 | 0.65 | -8.336 | -29.1 | 30.2774 | 18:51
19 | 190 | 0 | 0 | 0.65 | -8.395 | -29.96 | 31.1207 | 18:52
20 | 200 | 0 | 0 | 0.64 | -8.416 | -30.62 | 31.762 | 18:52
21 | 210 | 0 | 0 | 0.66 | -8.41 | -31.3 | 32.4169 | 18:52
22 | 220 | 0 | 0 | 0.66 | -8.404 | -32.28 | 33.3626 | 18:53
23 | 230 | 0 | 0 | 0.68 | -8.392 | -33.6 | 34.6388 | 18:53
24 | 240 | 0 | 0 | 0.66 | -8.407 | -34.81 | 35.8169 | 18:53
25 | 250 | 0 | 0 | 0.68 | -8.44 | -35.48 | 36.4764 | 18:54
26 | 260 | 0 | 0 | 0.66 | -8.489 | -35.13 | 36.1471 | 18:54
27 | 270 | 0 | 0 | 0.62 | -8.556 | -32.76 | 33.8645 | 18:54
28 | 280 | 0 | 0 | 0.55 | -8.652 | -27.56 | 28.8914 | 18:55
29 | 290 | 0 | 0 | 0.47 | -8.747 | -19.38 | 21.2677 | 18:55
30 | 300 | 0 | 0 | 0.34 | -8.774 | -8.92 | 12.5166 | 18:55
31 | 310 | 0 | 0 | 0.22 | -8.721 | 2.72 | 9.138 | 18:56
32 | 320 | 0 | 0 | 0.11 | -8.569 | 13.43 | 15.9313 | 18:56
33 | 330 | 0 | 0 | 0.07 | -8.335 | 21.69 | 23.2365 | 18:56
34 | 340 | 0 | 0 | 0.1 | -8.103 | 26.1 | 27.3291 | 18:57
35 | 350 | 0 | 0 | 0.15 | -7.875 | 26.28 | 27.435 | 18:57
36 | 360 | 0 | 0 | 0.28 | -7.713 | 22.43 | 23.7207 | 18:57
37 | 370 | 0 | 0 | 0.42 | -7.658 | 15.76 | 17.5271 | 18:58
38 | 380 | 0 | 0 | 0.53 | -7.698 | 8.11 | 11.1943 | 18:58
39 | 390 | 0 | 0 | 0.62 | -7.811 | 0.97 | 7.8954 | 18:58
40 | 400 | 0 | 0 | 0.66 | -7.934 | -4.79 | 9.2913 | 18:59
41 | 0 | 0 | 0 | 0.57 | -8.388 | -30.18 | 31.3292 | 18:59
41 | 0 | 0 | 0 | 0.4 | -8.317 | -8 | 11.547 | 18:59

## 14285right
XYZ MAPPING TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Wed, Feb 04, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
LOCKED AT 6.7T
Magnet locked in for 14+ hours
scanning down center of magnet, offset t the right 5mm (as see from side opposite of red vacuum plug)
| Z-axis | Vertical | Horizontal
Step | Z (mm) | Y (mm) | X (mm) | Field (Tesla) | Status | Time
0 | -29 | -0.1 | 5 | 6.6677 | **N0** | 11:28
0 | -28 | -0.1 | 5 | 6.6981 | LOCK | 11:29
0 | -27 | -0.1 | 5 | 6.6982 | LOCK | 11:29
0 | -26 | -0.1 | 5 | 6.6984 | LOCK | 11:30
0 | -25 | -0.1 | 5 | 6.6985 | LOCK | 11:30
0 | -25 | -0.1 | 5 | 6.6985 | LOCK | 11:30
0 | -24 | -0.1 | 5 | 6.6986 | LOCK | 11:30
0 | -23 | -0.1 | 5 | 6.6988 | LOCK | 11:30
0 | -22 | -0.1 | 5 | 6.6989 | LOCK | 11:32
0 | -21 | -0.1 | 5 | 6.699 | LOCK | 11:33
0 | -20 | -0.1 | 5 | 6.6991 | LOCK | 11:33
0 | -19 | -0.1 | 5 | 6.6992 | LOCK | 11:33
0 | -18 | -0.1 | 5 | 6.6993 | LOCK | 11:34
0 | -17 | -0.1 | 5 | 6.6993 | LOCK | 11:34
0 | -16 | -0.1 | 5 | 6.6994 | LOCK | 11:34
0 | -15 | -0.1 | 5 | 6.6995 | LOCK | 11:34
0 | -14 | -0.1 | 5 | 6.6996 | LOCK | 11:34
0 | -13 | -0.1 | 5 | 6.6996 | LOCK | 11:35
0 | -12 | -0.1 | 5 | 6.6997 | LOCK | 11:35
0 | -11 | -0.1 | 5 | 6.6998 | LOCK | 11:35
0 | -10 | -0.1 | 5 | 6.6998 | LOCK | 11:35
0 | -9 | -0.1 | 5 | 6.6999 | LOCK | 11:36
0 | -8 | -0.1 | 5 | 6.6999 | LOCK | 11:36
0 | -7 | -0.1 | 5 | 6.7 | LOCK | 11:36
0 | -6 | -0.1 | 5 | 6.7 | LOCK | 11:36
0 | -5 | -0.1 | 5 | 6.7 | LOCK | 11:36
0 | -4 | -0.1 | 5 | 6.7 | LOCK | 11:37
0 | -3 | -0.1 | 5 | 6.7001 | LOCK | 11:37
0 | -2 | -0.1 | 5 | 6.7001 | LOCK | 11:37
0 | -1 | -0.1 | 5 | 6.7001 | LOCK | 11:37
0 | 0 | -0.1 | 5 | 6.7001 | LOCK | 11:38
0 | 1 | -0.1 | 5 | 6.7001 | LOCK | 11:38
0 | 2 | -0.1 | 5 | 6.7001 | LOCK | 11:38
0 | 3 | -0.1 | 5 | 6.7001 | LOCK | 11:38
0 | 4 | -0.1 | 5 | 6.7001 | LOCK | 11:39
0 | 5 | -0.1 | 5 | 6.7001 | LOCK | 11:39
0 | 6 | -0.1 | 5 | 6.7001 | LOCK | 11:39
0 | 7 | -0.1 | 5 | 6.7001 | LOCK | 11:39
0 | 8 | -0.1 | 5 | 6.7001 | LOCK | 11:40
0 | 9 | -0.1 | 5 | 6.7 | LOCK | 11:40
0 | 10 | -0.1 | 5 | 6.7 | LOCK | 11:40
0 | 11 | -0.1 | 5 | 6.7 | LOCK | 11:40
0 | 12 | -0.1 | 5 | 6.7 | LOCK | 11:40
0 | 13 | -0.1 | 5 | 6.6999 | LOCK | 11:41
0 | 14 | -0.1 | 5 | 6.6999 | LOCK | 11:41
0 | 15 | -0.1 | 5 | 6.6998 | LOCK | 11:41
0 | 16 | -0.1 | 5 | 6.6998 | LOCK | 11:41
0 | 17 | -0.1 | 5 | 6.6997 | LOCK | 11:42
0 | 18 | -0.1 | 5 | 6.6997 | LOCK | 11:42
0 | 19 | -0.1 | 5 | 6.6996 | LOCK | 11:42
0 | 20 | -0.1 | 5 | 6.6995 | LOCK | 11:42
0 | 21 | -0.1 | 5 | 6.6995 | LOCK | 11:43
0 | 22 | -0.1 | 5 | 6.6994 | LOCK | 11:43
0 | 23 | -0.1 | 5 | 6.6993 | LOCK | 11:43
0 | 24 | -0.1 | 5 | 6.6992 | LOCK | 11:43
0 | 1 | 0 | 0 | 6.6991 | LOCK | 11:44
0 | 2 | 0 | 0 | 6.699 | LOCK | 11:45
0 | 3 | 0 | 0 | 6.6989 | LOCK | 11:45
0 | 4 | 0 | 0 | 6.6988 | LOCK | 11:45
0 | 5 | 0 | 0 | 6.6987 | LOCK | 11:46
0 | 6 | 0 | 0 | 6.6986 | LOCK | 11:46
0 | 7 | 0 | 0 | 6.6985 | LOCK | 11:46
0 | 8 | 0 | 0 | 6.7035 | SIGNAL | 11:46

## 14285left
XYZ MAPPING TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Wed, Feb 04, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
LOCKED AT 6.7T
Magnet locked in for 14+ hours
scanning down center of magnet, offset t the left (as see from side opposite of red vacuum plug)
| Z-axis | Vertical | Horizontal
Step | Z (mm) | Y (mm) | X (mm) | Field (Tesla) | Status | Time
0 | -28 | -0.1 | -5 | 6.6982 | LOCK | 11:03
0 | -27 | -0.1 | -5 | 6.6983 | LOCK | 11:04
0 | -26 | -0.1 | -5 | 6.6984 | SIGNAL | 11:04
0 | -25 | -0.1 | -5 | 6.6986 | LOCK | 11:04
0 | -24 | -0.1 | -5 | 6.6987 | LOCK | 11:05
0 | -23 | -0.1 | -5 | 6.6988 | LOCK | 11:05
0 | -22 | -0.1 | -5 | 6.6989 | LOCK | 11:05
0 | -21 | -0.1 | -5 | 6.699 | LOCK | 11:05
0 | -20 | -0.1 | -5 | 6.6991 | LOCK | 11:06
0 | -19 | -0.1 | -5 | 6.6992 | LOCK | 11:06
0 | -18 | -0.1 | -5 | 6.6993 | LOCK | 11:06
0 | -17 | -0.1 | -5 | 6.6994 | LOCK | 11:07
0 | -16 | -0.1 | -5 | 6.6995 | LOCK | 11:07
0 | -15 | -0.1 | -5 | 6.6995 | LOCK | 11:07
0 | -14 | -0.1 | -5 | 6.6996 | LOCK | 11:07
0 | -13 | -0.1 | -5 | 6.6997 | LOCK | 11:08
0 | -12 | -0.1 | -5 | 6.6997 | LOCK | 11:08
0 | -11 | -0.1 | -5 | 6.6998 | LOCK | 11:08
0 | -10 | -0.1 | -5 | 6.6998 | LOCK | 11:08
0 | -9 | -0.1 | -5 | 6.6999 | LOCK | 11:09
0 | -8 | -0.1 | -5 | 6.6999 | LOCK | 11:09
0 | -7 | -0.1 | -5 | 6.6999 | LOCK | 11:09
0 | -6 | -0.1 | -5 | 6.7 | LOCK | 11:09
0 | -5 | -0.1 | -5 | 6.7 | LOCK | 11:10
0 | -4 | -0.1 | -5 | 6.7 | LOCK | 11:10
0 | -3 | -0.1 | -5 | 6.7 | LOCK | 11:10
0 | -2 | -0.1 | -5 | 6.7001 | LOCK | 11:10
0 | -1 | -0.1 | -5 | 6.7001 | LOCK | 11:11
0 | 0 | -0.1 | -5 | 6.7001 | LOCK | 11:11
0 | 1 | -0.1 | -5 | 6.7001 | LOCK | 11:11
0 | 2 | -0.1 | -5 | 6.7001 | LOCK | 11:11
0 | 3 | -0.1 | -5 | 6.7001 | LOCK | 11:12
0 | 4 | -0.1 | -5 | 6.7001 | LOCK | 11:12
0 | 5 | -0.1 | -5 | 6.7001 | LOCK | 11:12
0 | 6 | -0.1 | -5 | 6.7 | LOCK | 11:12
0 | 7 | -0.1 | -5 | 6.7 | LOCK | 11:13
0 | 8 | -0.1 | -5 | 6.7 | LOCK | 11:13
0 | 9 | -0.1 | -5 | 6.7 | LOCK | 11:13
0 | 10 | -0.1 | -5 | 6.7 | LOCK | 11:13
0 | 11 | -0.1 | -5 | 6.6999 | LOCK | 11:14
0 | 12 | -0.1 | -5 | 6.6999 | LOCK | 11:14
0 | 13 | -0.1 | -5 | 6.6998 | LOCK | 11:14
0 | 14 | -0.1 | -5 | 6.6998 | LOCK | 11:14
0 | 15 | -0.1 | -5 | 6.6998 | LOCK | 11:15
0 | 16 | -0.1 | -5 | 6.6997 | LOCK | 11:15
0 | 17 | -0.1 | -5 | 6.6996 | LOCK | 11:15
0 | 18 | -0.1 | -5 | 6.6996 | LOCK | 11:15
0 | 19 | -0.1 | -5 | 6.6995 | LOCK | 11:16
0 | 20 | -0.1 | -5 | 6.6995 | LOCK | 11:16
0 | 21 | -0.1 | -5 | 6.6994 | LOCK | 11:16
0 | 22 | -0.1 | -5 | 6.6993 | LOCK | 11:16
0 | 23 | -0.1 | -5 | 6.6992 | LOCK | 11:17
0 | 24 | -0.1 | -5 | 6.6991 | LOCK | 11:17
0 | 25 | -0.1 | -5 | 6.6991 | LOCK | 11:17
0 | 26 | -0.1 | -5 | 6.699 | LOCK | 11:17
0 | 27 | -0.1 | -5 | 6.6989 | LOCK | 11:18
0 | 28 | -0.1 | -5 | 6.6988 | LOCK | 11:18
0 | 29 | -0.1 | -5 | 6.6987 | LOCK | 11:18
0 | 30 | -0.1 | -5 | 6.6985 | LOCK | 11:18
0 | 31 | -0.1 | -5 | 6.6984 | LOCK | 11:19
0 | 32 | -0.1 | -5 | 6.6983 | LOCK | 11:19
0 | 33 | -0.1 | -5 | 6.6982 | LOCK | 11:19

## 6.7T 5mm below
XYZ MAPPING TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Wed, Feb 04, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
LOCKED AT 6.7T
for 14 hours
SECOND PART, TAKESOVER AT -21MM OF THE BOTTOM AXIS SCAN
Probe is ~5mm below center
| Z-axis | Vertical | Horizontal
Step | Z (mm) | Y (mm) | X (mm) | Field (Tesla) | Status | Time
88 | -28 | -5.1 | 0 | 6.6982 | LOCK | 9:17
88 | -27 | -5.1 | 0 | 6.6983 | LOCK | 9:18
88 | -26 | -5.1 | 0 | 6.6985 | LOCK | 9:18
88 | -25 | -5.1 | 0 | 6.6986 | LOCK | 9:19
88 | -25 | -5.1 | 0 | 6.6986 | LOCK | 9:19
88 | -24 | -5.1 | 0 | 6.6987 | LOCK | 9:20
88 | -23 | -5.1 | 0 | 6.6988 | LOCK | 9:20
88 | -22 | -5.1 | 0 | 6.699 | LOCK | 9:25
0 | -21 | -5.1 | 0 | 6.6991 | LOCK | 9:28
0 | -20 | -5.1 | 0 | 6.6992 | LOCK | 9:29
0 | -19 | -5.1 | 0 | 6.6993 | LOCK | 9:30
0 | -18 | -5.1 | 0 | 6.6994 | LOCK | 9:30
0 | -18 | -5.1 | 0 | 6.6994 | LOCK | 9:30
0 | -17 | -5.1 | 0 | 6.6994 | LOCK | 9:31
0 | -16 | -5.1 | 0 | 6.6995 | LOCK | 9:31
0 | -15 | -5.1 | 0 | 6.6996 | LOCK | 9:32
0 | -14 | -5.1 | 0 | 6.6997 | LOCK | 9:32
0 | -13 | -5.1 | 0 | 6.6997 | LOCK | 9:32
0 | -12 | -5.1 | 0 | 6.6998 | LOCK | 9:32
0 | -11 | -5.1 | 0 | 6.6999 | LOCK | 9:33
0 | -10 | -5.1 | 0 | 6.6999 | LOCK | 9:33
0 | -9 | -5.1 | 0 | 6.6999 | LOCK | 9:33
0 | -8 | -5.1 | 0 | 6.7 | LOCK | 9:35
0 | -7 | -5.1 | 0 | 6.7 | LOCK | 9:36
0 | -6 | -5.1 | 0 | 6.7001 | LOCK | 9:37
0 | -5 | -5.1 | 0 | 6.7001 | LOCK | 9:37
0 | -4 | -5.1 | 0 | 6.7001 | LOCK | 9:38
0 | -3 | -5.1 | 0 | 6.7001 | LOCK | 9:39
0 | -2 | -5.1 | 0 | 6.7002 | LOCK | 9:39
0 | -1 | -5.1 | 0 | 6.7002 | LOCK | 9:39
0 | 0 | -5.1 | 0 | 6.7002 | LOCK | 9:40
0 | 1 | -5.1 | 0 | 6.7002 | LOCK | 9:40
0 | 2 | -5.1 | 0 | 6.7002 | LOCK | 9:41
0 | 2 | -5.1 | 0 | 6.7002 | LOCK | 9:41
0 | 3 | -5.1 | 0 | 6.7002 | LOCK | 9:42
0 | 4 | -5.1 | 0 | 6.7002 | LOCK | 9:42
0 | 5 | -5.1 | 0 | 6.7002 | LOCK | 9:43
0 | 6 | -5.1 | 0 | 6.7002 | LOCK | 9:43
0 | 7 | -5.1 | 0 | 6.7001 | LOCK | 9:44
0 | 8 | -5.1 | 0 | 6.7001 | LOCK | 9:44
0 | 9 | -5.1 | 0 | 6.7001 | LOCK | 9:46
0 | 10 | -5.1 | 0 | 6.7001 | LOCK | 9:46
0 | 11 | -5.1 | 0 | 6.7 | LOCK | 9:49
0 | 12 | -5.1 | 0 | 6.7 | LOCK | 9:49
0 | 13 | -5.1 | 0 | 6.7 | LOCK | 9:50
0 | 14 | -5.1 | 0 | 6.6999 | LOCK | 9:50
0 | 15 | -5.1 | 0 | 6.6999 | LOCK | 9:50
0 | 16 | -5.1 | 0 | 6.6998 | LOCK | 9:51
0 | 17 | -5.1 | 0 | 6.6998 | LOCK | 9:51
0 | 18 | -5.1 | 0 | 6.6997 | LOCK | 9:51
0 | 19 | -5.1 | 0 | 6.6996 | LOCK | 9:52
0 | 20 | -5.1 | 0 | 6.6996 | LOCK | 9:52
0 | 20 | -5.1 | 0 | 6.6996 | LOCK | 9:52
0 | 21 | -5.1 | 0 | 6.6995 | LOCK | 9:52
0 | 22 | -5.1 | 0 | 6.6994 | LOCK | 9:53
0 | 23 | -5.1 | 0 | 6.6993 | LOCK | 9:53
0 | 24 | -5.1 | 0 | 6.6993 | LOCK | 9:53
0 | 25 | -5.1 | 0 | 6.6992 | LOCK | 9:54
0 | 25 | -5.1 | 0 | 6.6992 | LOCK | 9:54
0 | 26 | -5.1 | 0 | 6.6991 | LOCK | 9:54
0 | 27 | -5.1 | 0 | 6.699 | LOCK | 9:54
0 | 28 | -5.1 | 0 | 6.6989 | LOCK | 9:55
0 | 29 | -5.1 | 0 | 6.6988 | LOCK | 9:55
0 | 30 | -5.1 | 0 | 6.6988 | LOCK | 9:55
0 | 30 | -5.1 | 0 | 6.6987 | LOCK | 9:55
0 | 31 | -5.1 | 0 | 6.6986 | LOCK | 9:55
0 | 32 | -5.1 | 0 | 6.6984 | LOCK | 9:56
0 | 33 | -5.1 | 0 | 6.6983 | LOCK | 9:56
0 | 34 | -5.1 | 0 | 6.7056 | **N0** | 9:56
0 | 35 | -5.1 | 0 | 6.726 | **N0** | 9:57
0 | 36 | -5.1 | 0 | 6.6387 | **N0** | 9:57
0 | 37 | -5.1 | 0 | 6.6394 | **N0** | 9:57
0 | 38 | -5.1 | 0 | 6.7246 | **N0** | 9:58
0 | 39 | -5.1 | 0 | 6.7214 | **N0** | 9:58

## 6.7T 5mm above
XYZ MAPPING TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Wed, Feb 04, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
LOCKED AT 6.7T
for 14 hours
Probe is ~5mm above center
| Z-axis | Vertical | Horizontal
Step | Z (mm) | Y (mm) | X (mm) | Field (Tesla) | Status | Time
88 | -27 | 4.9 | 0 | 6.6982 | LOCK | 8:46
88 | -26 | 4.9 | 0 | 6.6983 | LOCK | 8:49
88 | -25 | 4.9 | 0 | 6.6984 | LOCK | 8:49
88 | -24 | 4.9 | 0 | 6.6985 | LOCK | 8:49
88 | -23 | 4.9 | 0 | 6.6987 | LOCK | 8:50
88 | -22 | 4.9 | 0 | 6.6988 | LOCK | 8:50
88 | -21 | 4.9 | 0 | 6.6989 | LOCK | 8:51
88 | -20 | 4.9 | 0 | 6.699 | LOCK | 8:51
88 | -19 | 4.9 | 0 | 6.6991 | LOCK | 8:51
88 | -18 | 4.9 | 0 | 6.6992 | LOCK | 8:52
88 | -17 | 4.9 | 0 | 6.6993 | LOCK | 8:52
88 | -16 | 4.9 | 0 | 6.6994 | LOCK | 8:53
88 | -15 | 4.9 | 0 | 6.6994 | LOCK | 8:53
88 | -14 | 4.9 | 0 | 6.6995 | LOCK | 8:53
88 | -13 | 4.9 | 0 | 6.6996 | LOCK | 8:54
88 | -12 | 4.9 | 0 | 6.6996 | LOCK | 8:54
88 | -11 | 4.9 | 0 | 6.6997 | LOCK | 8:54
88 | -10 | 4.9 | 0 | 6.6997 | LOCK | 8:54
88 | -9 | 4.9 | 0 | 6.6998 | LOCK | 8:55
88 | -9 | 4.9 | 0 | 6.6998 | LOCK | 8:55
88 | -8 | 4.9 | 0 | 6.6998 | SIGNAL | 8:55
88 | -7 | 4.9 | 0 | 6.6999 | LOCK | 8:56
88 | -6 | 4.9 | 0 | 6.6999 | LOCK | 8:56
88 | -5 | 4.9 | 0 | 6.6999 | LOCK | 8:56
88 | -4 | 4.9 | 0 | 6.6999 | LOCK | 8:57
88 | -3 | 4.9 | 0 | 6.7 | LOCK | 8:58
88 | -2 | 4.9 | 0 | 6.7 | LOCK | 8:58
88 | -1 | 4.9 | 0 | 6.7 | LOCK | 8:58
88 | 0 | 4.9 | 0 | 6.7 | LOCK | 8:59
88 | 1 | 4.9 | 0 | 6.7 | LOCK | 8:59
88 | 2 | 4.9 | 0 | 6.7 | LOCK | 8:59
88 | 3 | 4.9 | 0 | 6.7 | LOCK | 8:59
88 | 4 | 4.9 | 0 | 6.7 | LOCK | 9:00
88 | 5 | 4.9 | 0 | 6.7 | LOCK | 9:00
88 | 6 | 4.9 | 0 | 6.7 | LOCK | 9:00
88 | 7 | 4.9 | 0 | 6.7 | LOCK | 9:01
88 | 8 | 4.9 | 0 | 6.7 | LOCK | 9:02
88 | 9 | 4.9 | 0 | 6.6999 | LOCK | 9:02
88 | 10 | 4.9 | 0 | 6.6999 | LOCK | 9:02
88 | 11 | 4.9 | 0 | 6.6999 | LOCK | 9:02
88 | 12 | 4.9 | 0 | 6.6998 | LOCK | 9:02
88 | 13 | 4.9 | 0 | 6.6998 | LOCK | 9:03
88 | 14 | 4.9 | 0 | 6.6998 | LOCK | 9:03
88 | 15 | 4.9 | 0 | 6.6997 | LOCK | 9:03
88 | 16 | 4.9 | 0 | 6.6997 | LOCK | 9:04
88 | 17 | 4.9 | 0 | 6.6996 | LOCK | 9:04
88 | 18 | 4.9 | 0 | 6.6996 | LOCK | 9:05
88 | 19 | 4.9 | 0 | 6.6995 | LOCK | 9:05
88 | 20 | 4.9 | 0 | 6.6994 | LOCK | 9:05
88 | 21 | 4.9 | 0 | 6.6994 | LOCK | 9:05
88 | 22 | 4.9 | 0 | 6.6993 | LOCK | 9:06
88 | 23 | 4.9 | 0 | 6.6992 | LOCK | 9:06
88 | 24 | 4.9 | 0 | 6.6991 | LOCK | 9:06
88 | 25 | 4.9 | 0 | 6.699 | LOCK | 9:07
88 | 26 | 4.9 | 0 | 6.6989 | LOCK | 9:07
88 | 27 | 4.9 | 0 | 6.6988 | LOCK | 9:08
88 | 28 | 4.9 | 0 | 6.6987 | LOCK | 9:08
88 | 29 | 4.9 | 0 | 6.6986 | LOCK | 9:08
88 | 30 | 4.9 | 0 | 6.6985 | LOCK | 9:09
88 | 31 | 4.9 | 0 | 6.6984 | LOCK | 9:09
88 | 32 | 4.9 | 0 | 6.6982 | LOCK | 9:09
88 | 33 | 4.9 | 0 | 6.6981 | LOCK | 9:10
88 | 34 | 4.9 | 0 | 6.698 | LOCK | 9:10
88 | 35 | 4.9 | 0 | 6.6978 | LOCK | 9:11
88 | 36 | 4.9 | 0 | 6.6976 | LOCK | 9:11
Formula for Bz which fits all the data.
Bz = -2E-06*z2 + 1E-05*z*(1-(.1/8)*((x/2.5)2-2*(x/2.5)) + 6.7001-(.0001/5)*y
Worst (& dominant) gradient is in the y direction.
Center in x direction is at 2.5mm (with quadractic gradient)
Cneter in z direction is at 2.5mm (with quadratic gradient)
Center in y direction is at 0mm (but with linear gradient) .. this term totally dominates the inhomogeneity

## 6.7T Scan after locked over nig
This data was collected manually which is why "Step" is 88 for all.
XYZ MAPPING TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Wed, Feb 04, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
LOCKED AT 6.7T
for 14 hours
| Z-axis | Vertical | Horizontal
Step | Z (mm) | Y (mm) | X (mm) | Field (Tesla) | Status | Time
88 | -26 | -0.1 | 0 | 6.6983 | LOCK | 8:09
88 | -26 | -0.1 | 0 | 6.6983 | LOCK | 8:10
88 | -26 | -0.1 | 0 | 6.6983 | LOCK | 8:10
88 | -26 | -0.1 | 0 | 6.6983 | LOCK | 8:10
88 | -26 | -0.1 | 0 | 6.6983 | LOCK | 8:11
88 | -26 | -0.1 | 0 | 6.6983 | LOCK | 8:11
88 | -25.9 | -0.1 | 0 | 6.6983 | LOCK | 8:11
88 | -25.9 | -0.1 | 0 | 6.6984 | LOCK | 8:11
88 | -25.9 | -0.1 | 0 | 6.6984 | LOCK | 8:12
88 | -25.9 | -0.1 | 0 | 6.6984 | LOCK | 8:12
88 | -25.8 | -0.1 | 0 | 6.6984 | LOCK | 8:12
88 | -25.7 | -0.1 | 0 | 6.6984 | LOCK | 8:13
88 | -25.6 | -0.1 | 0 | 6.6984 | LOCK | 8:13
88 | -25.5 | -0.1 | 0 | 6.6984 | LOCK | 8:14
88 | -25.4 | -0.1 | 0 | 6.6984 | LOCK | 8:14
88 | -25.3 | -0.1 | 0 | 6.6984 | LOCK | 8:14
88 | -25.2 | -0.1 | 0 | 6.6984 | LOCK | 8:15
88 | -25.1 | -0.1 | 0 | 6.6985 | LOCK | 8:16
88 | -25 | -0.1 | 0 | 6.6985 | LOCK | 8:16
88 | -24.9 | -0.1 | 0 | 6.6985 | LOCK | 8:16
88 | -24.8 | -0.1 | 0 | 6.6985 | LOCK | 8:17
88 | -24.7 | -0.1 | 0 | 6.6985 | LOCK | 8:17
88 | -24.6 | -0.1 | 0 | 6.6985 | LOCK | 8:17
88 | -24.5 | -0.1 | 0 | 6.6985 | LOCK | 8:18
88 | -24.4 | -0.1 | 0 | 6.6985 | LOCK | 8:22
88 | -24 | -0.1 | 0 | 6.6986 | LOCK | 8:23
88 | -24 | -0.1 | 0 | 6.6986 | LOCK | 8:23
88 | -23 | -0.1 | 0 | 6.6987 | LOCK | 8:24
88 | -22 | -0.1 | 0 | 6.6988 | LOCK | 8:24
88 | -21 | -0.1 | 0 | 6.699 | LOCK | 8:24
88 | -20 | -0.1 | 0 | 6.6991 | LOCK | 8:25
88 | -19 | -0.1 | 0 | 6.6992 | LOCK | 8:25
88 | -18 | -0.1 | 0 | 6.6992 | LOCK | 8:25
88 | -17 | -0.1 | 0 | 6.6993 | LOCK | 8:26
88 | -16 | -0.1 | 0 | 6.6994 | LOCK | 8:26
88 | -15 | -0.1 | 0 | 6.6995 | LOCK | 8:27
88 | -14 | -0.1 | 0 | 6.6996 | LOCK | 8:27
88 | -13 | -0.1 | 0 | 6.6996 | LOCK | 8:27
88 | -12 | -0.1 | 0 | 6.6997 | LOCK | 8:28
88 | -11 | -0.1 | 0 | 6.6997 | LOCK | 8:28
88 | -10 | -0.1 | 0 | 6.6998 | LOCK | 8:28
88 | -9 | -0.1 | 0 | 6.6998 | LOCK | 8:29
88 | -8 | -0.1 | 0 | 6.6999 | LOCK | 8:29
88 | -7 | -0.1 | 0 | 6.6999 | LOCK | 8:29
88 | -6 | -0.1 | 0 | 6.6999 | LOCK | 8:30
88 | -5 | -0.1 | 0 | 6.7 | LOCK | 8:30
88 | -4 | -0.1 | 0 | 6.7 | LOCK | 8:30
88 | -3 | -0.1 | 0 | 6.7 | LOCK | 8:30
88 | -2 | -0.1 | 0 | 6.7 | LOCK | 8:31
88 | -1 | -0.1 | 0 | 6.7001 | LOCK | 8:31
88 | 0 | -0.1 | 0 | 6.7001 | LOCK | 8:31
88 | 1 | -0.1 | 0 | 6.7001 | LOCK | 8:31
88 | 2 | -0.1 | 0 | 6.7001 | LOCK | 8:32
88 | 3 | -0.1 | 0 | 6.7001 | LOCK | 8:32
88 | 4 | -0.1 | 0 | 6.7001 | LOCK | 8:32
88 | 5 | -0.1 | 0 | 6.7001 | LOCK | 8:32
88 | 6 | -0.1 | 0 | 6.7001 | LOCK | 8:33
88 | 7 | -0.1 | 0 | 6.7 | LOCK | 8:33
88 | 8 | -0.1 | 0 | 6.7 | LOCK | 8:33
88 | 9 | -0.1 | 0 | 6.7 | LOCK | 8:33
88 | 10 | -0.1 | 0 | 6.7 | LOCK | 8:34
88 | 11 | -0.1 | 0 | 6.6999 | LOCK | 8:34
88 | 12 | -0.1 | 0 | 6.6999 | LOCK | 8:35
88 | 13 | -0.1 | 0 | 6.6999 | LOCK | 8:35
88 | 14 | -0.1 | 0 | 6.6998 | LOCK | 8:35
88 | 15 | -0.1 | 0 | 6.6998 | LOCK | 8:35
88 | 16 | -0.1 | 0 | 6.6997 | LOCK | 8:36
88 | 17 | -0.1 | 0 | 6.6997 | LOCK | 8:36
88 | 18 | -0.1 | 0 | 6.6996 | LOCK | 8:36
88 | 19 | -0.1 | 0 | 6.6995 | LOCK | 8:37
88 | 20 | -0.1 | 0 | 6.6995 | LOCK | 8:37
88 | 21 | -0.1 | 0 | 6.6994 | LOCK | 8:37
88 | 22 | -0.1 | 0 | 6.6993 | LOCK | 8:37
88 | 23 | -0.1 | 0 | 6.6992 | LOCK | 8:38
88 | 24 | -0.1 | 0 | 6.6992 | LOCK | 8:38
88 | 25 | -0.1 | 0 | 6.6991 | LOCK | 8:38
88 | 26 | -0.1 | 0 | 6.699 | LOCK | 8:38
88 | 27 | -0.1 | 0 | 6.6989 | LOCK | 8:39
88 | 28 | -0.1 | 0 | 6.6988 | LOCK | 8:39
88 | 29 | -0.1 | 0 | 6.6986 | LOCK | 8:39
88 | 30 | -0.1 | 0 | 6.6985 | LOCK | 8:40
88 | 31 | -0.1 | 0 | 6.6984 | LOCK | 8:40
88 | 32 | -0.1 | 0 | 6.6983 | LOCK | 8:40
88 | 33 | -0.1 | 0 | 6.6981 | LOCK | 8:41
88 | 34 | -0.1 | 0 | 6.698 | LOCK | 8:41
88 | 35 | -0.1 | 0 | 6.6978 | LOCK | 8:41
y=-2E-06*(x-2.5)^2+6.7001
center field is Z=2.5mm position
Note on E67: Ronny Efferson: Call this Center point +/-0.001% for +/-3mm on this axis.

## 14285Drift
We had to move the probe from 1mm to 2mm so it was at different locations, otherwise the unit would not collect the data.
XYZ MAPPING TABLE UNIFORMITY DATA FILE
Test Engineer: Scott Choroba
Date: Tues, Feb 03, 2015
Magnet Model: 14285 TRIUMF
Magnet Serial No:14285
LOCKED AT 6.5T
| Z-axis | Vertical | Horizontal
Step | Z (mm) | Y (mm) | X (mm) | Field (Tesla) | Status | Time | Time Dec.
4 | 2 | 0 | 0 | 6.7002 | LOCK | 18:59 | 18.9775
5 | 1 | 0 | 0 | 6.7002 | LOCK | 19:08 | 19.1317
6 | 2 | 0 | 0 | 6.7002 | LOCK | 19:17 | 19.2858
7 | 1 | 0 | 0 | 6.7002 | LOCK | 19:26 | 19.44
8 | 2 | 0 | 0 | 6.7002 | LOCK | 19:36 | 19.5942
9 | 1 | 0 | 0 | 6.7001 | LOCK | 19:45 | 19.7483
10 | 2 | 0 | 0 | 6.7001 | LOCK | 19:54 | 19.9025
11 | 1 | 0 | 0 | 6.7001 | LOCK | 20:03 | 20.0567
12 | 2 | 0 | 0 | 6.7001 | LOCK | 20:13 | 20.2108
13 | 1 | 0 | 0 | 6.7001 | LOCK | 20:22 | 20.365
14 | 2 | 0 | 0 | 6.7001 | LOCK | 20:31 | 20.5192
15 | 1 | 0 | 0 | 6.7001 | LOCK | 20:40 | 20.6733
16 | 2 | 0 | 0 | 6.7001 | LOCK | 20:50 | 20.8275
17 | 1 | 0 | 0 | 6.7001 | LOCK | 20:59 | 20.9817
18 | 2 | 0 | 0 | 6.7001 | LOCK | 21:08 | 21.1358
19 | 1 | 0 | 0 | 6.7001 | LOCK | 21:17 | 21.29
20 | 2 | 0 | 0 | 6.7001 | LOCK | 21:27 | 21.4442
21 | 1 | 0 | 0 | 6.7001 | LOCK | 21:36 | 21.5983
22 | 2 | 0 | 0 | 6.7001 | LOCK | 21:45 | 21.7525
23 | 1 | 0 | 0 | 6.7001 | LOCK | 21:54 | 21.9067
24 | 2 | 0 | 0 | 6.7001 | LOCK | 22:04 | 22.0608
25 | 1 | 0 | 0 | 6.7001 | LOCK | 22:13 | 22.215
26 | 2 | 0 | 0 | 6.7001 | LOCK | 22:22 | 22.3692
27 | 1 | 0 | 0 | 6.7001 | LOCK | 22:31 | 22.5233
28 | 2 | 0 | 0 | 6.7001 | LOCK | 22:41 | 22.6775
29 | 1 | 0 | 0 | 6.7001 | LOCK | 22:50 | 22.8317
30 | 2 | 0 | 0 | 6.7001 | LOCK | 22:59 | 22.9858
31 | 1 | 0 | 0 | 6.7001 | LOCK | 23:08 | 23.14
32 | 2 | 0 | 0 | 6.7001 | LOCK | 23:18 | 23.2942
33 | 1 | 0 | 0 | 6.7001 | LOCK | 23:27 | 23.4483
34 | 2 | 0 | 0 | 6.7001 | LOCK | 23:36 | 23.6025
35 | 1 | 0 | 0 | 6.7001 | LOCK | 23:45 | 23.7567
36 | 2 | 0 | 0 | 6.7001 | LOCK | 23:55 | 23.9108
37 | 1 | 0 | 0 | 6.7001 | LOCK | 24:04 | 24.065
38 | 2 | 0 | 0 | 6.7001 | LOCK | 24:13 | 24.2192
39 | 1 | 0 | 0 | 6.7001 | LOCK | 24:22 | 24.3733
40 | 2 | 0 | 0 | 6.7001 | LOCK | 24:32 | 24.5275
41 | 1 | 0 | 0 | 6.7001 | LOCK | 24:41 | 24.6817
42 | 2 | 0 | 0 | 6.7001 | LOCK | 24:50 | 24.8358
43 | 1 | 0 | 0 | 6.7001 | LOCK | 24:59 | 24.99
44 | 2 | 0 | 0 | 6.7001 | LOCK | 25:09 | 25.1442
45 | 1 | 0 | 0 | 6.7001 | LOCK | 25:18 | 25.2983
46 | 2 | 0 | 0 | 6.7001 | LOCK | 25:27 | 25.4525
47 | 1 | 0 | 0 | 6.7001 | LOCK | 25:36 | 25.6067
48 | 2 | 0 | 0 | 6.7001 | LOCK | 25:46 | 25.7608
49 | 1 | 0 | 0 | 6.7001 | LOCK | 25:55 | 25.915
50 | 2 | 0 | 0 | 6.7001 | LOCK | 26:04 | 26.0692
51 | 1 | 0 | 0 | 6.7001 | LOCK | 26:13 | 26.2233
52 | 2 | 0 | 0 | 6.7001 | LOCK | 26:23 | 26.3775
53 | 1 | 0 | 0 | 6.7001 | LOCK | 26:32 | 26.5317
54 | 2 | 0 | 0 | 6.7001 | LOCK | 26:41 | 26.6858
55 | 1 | 0 | 0 | 6.7001 | LOCK | 26:50 | 26.84
56 | 2 | 0 | 0 | 6.7001 | LOCK | 27:00 | 26.9942
57 | 1 | 0 | 0 | 6.7001 | LOCK | 27:09 | 27.1483
58 | 2 | 0 | 0 | 6.7001 | LOCK | 27:18 | 27.3025
59 | 1 | 0 | 0 | 6.7001 | LOCK | 27:27 | 27.4567
60 | 2 | 0 | 0 | 6.7001 | LOCK | 27:37 | 27.6108
61 | 1 | 0 | 0 | 6.7001 | LOCK | 27:46 | 27.765
62 | 2 | 0 | 0 | 6.7001 | LOCK | 27:55 | 27.9192
63 | 1 | 0 | 0 | 6.7001 | LOCK | 28:04 | 28.0733
64 | 2 | 0 | 0 | 6.7001 | LOCK | 28:14 | 28.2275
65 | 1 | 0 | 0 | 6.7001 | LOCK | 28:23 | 28.3817
66 | 2 | 0 | 0 | 6.7001 | LOCK | 28:32 | 28.5358
67 | 1 | 0 | 0 | 6.7001 | LOCK | 28:41 | 28.69
68 | 2 | 0 | 0 | 6.7001 | LOCK | 28:51 | 28.8442
69 | 1 | 0 | 0 | 6.7001 | LOCK | 29:00 | 28.9983
70 | 2 | 0 | 0 | 6.7001 | LOCK | 29:09 | 29.1525
71 | 1 | 0 | 0 | 6.7001 | LOCK | 29:18 | 29.3067
72 | 2 | 0 | 0 | 6.7001 | LOCK | 29:28 | 29.4608
73 | 1 | 0 | 0 | 6.7001 | LOCK | 29:37 | 29.615
74 | 2 | 0 | 0 | 6.7001 | LOCK | 29:46 | 29.7692
75 | 1 | 0 | 0 | 6.7001 | LOCK | 29:55 | 29.9233
76 | 2 | 0 | 0 | 6.7001 | LOCK | 30:05 | 30.0775
77 | 1 | 0 | 0 | 6.7001 | LOCK | 30:14 | 30.2317
78 | 2 | 0 | 0 | 6.7001 | LOCK | 30:23 | 30.3858
79 | 1 | 0 | 0 | 6.7001 | LOCK | 30:32 | 30.54
80 | 2 | 0 | 0 | 6.7001 | LOCK | 30:42 | 30.6942
81 | 1 | 0 | 0 | 6.7001 | LOCK | 30:51 | 30.8483
82 | 2 | 0 | 0 | 6.7001 | LOCK | 31:00 | 31.0025
83 | 1 | 0 | 0 | 6.7001 | LOCK | 31:09 | 31.1567
84 | 2 | 0 | 0 | 6.7001 | LOCK | 31:19 | 31.3108

## Homogeneity Summary
Formula for Bz which fits all the data.
Bz = -2E-06*z2 + 1E-05*z*(1-(.1/8)*((x/2.5)2-2*(x/2.5)) + 6.7001-(.0001/5)*y
Worst (& dominant) gradient is in the y direction.
Center in x direction is at 2.5mm (with quadractic gradient)
Cneter in z direction is at 2.5mm (with quadratic gradient)
Center in y direction is at 0mm (but with linear gradient) .. this term totally dominates the inhomogeneity